Insert the title of your
presentation here
Thomas Schietinger :: Paul Scherrer Institut
(with input from Andreas Lüdeke, Elmar Zehnder and others)
SLS Operation 2020
5th LAB Meeting, 6 November 2020

---
Overview SLS operation 2020
24/5 operation during LO-II/III periods
COVID-19 priority call
5 June:
restart of regular
user operations
First half of the year: limited operations from mid March
April shutdown cancelled!

---
Overview SLS operation 2020
Second half of the year: all user operation periods as initially planned

---
SLS operation statistics
(courtesy A. Lüdeke)
Year 2020 has been difficult:
  - Beam availability up to now 97.4%
  - Several weeks for COVID-19 research with no external users
  - Users send samples in, measurements by PSI staff
  - Mean time between distortions was relatively low, about 24 hours
data until 19-Oct-2020

---
Major issues
Very long interruptions (more than 5 hours downtime):
  - 31 January: Vacuum leak in linac RF waveguide – repair takes 33 hours.
  - 21 February: Arcs in linac RF structure after a power glitch; nominal power cannot be restored, must run at lower power. Vacuum window exchanged in shutdown, the first breaks again, the second holds – one day extension of shutdown.
  - 2 April: Water leak in cavity 2, temporary fix (5 hours downtime).
  - 17 June: Severe power glitch results in almost 8 hours of downtime; fortunately during beam development shifts; 3HC survived without problems.
  - 10/11 October (and again 20 October): Klystron problems in storage-ring RF station A4, broken connectors needed to be replaced. Several outages result in a total of 20 hours downtime. Problem still not completely solved, but reduced noise thanks to new cabeling and improved operational stability thanks to raised interlock thresholds.
Other problems (few hours downtime):
  - Water leak in A2 klystron, the Thales klystron was exchanged for an EEV klystron. Thales klystron repaired in the meantime.
  - 21 June: linac transmission problem – so far a mistery (still not solved and recurring occasionally!!). May be related to new LLRF system in linac.
  - 25 June: Fire in WSLA (building evacuation, 4 hours downtime) – see next slide.
  - 25 August: Problems with dipole power supply (cable lug, resistors) – about 4 hours downtime.
  - 29 August: Again dipole power supply problem, the actuator had to be exchanged (by crane) – 3 hours downtime.

---
Fire in WSLA (25 June)
- On 25 June 2020 a fire broke out in the X-ray bunker 45 (sector 29) of the SLS (TOMCAT beamline).
- The culprit was a standard propane based air conditioning device, which was operated to keep the temperature in the bunker bearable.
- Fortunately the concrete bunker walls prohibited further expansion of the fire – in other hutches the damage could have been much larger!
- The event triggered an
  event evaluation
  (AN-96-20-366)
  as well as an internal
  audit of the fire load
  along the SLS outer hull
  (AN-20-375).
- Strong recommendation:
  leave cooling and heating
  in your labs to the experts!
  (Sektion Gebäudetechnik)

---
Shutdown work and related activites
(courtesy E. Zehnder)
- New outside nitrogen gas tank – change to Linde Gas for supply
- Raising of hall ventilation and installation of service bridge at 5 m elevation
  - Temperature stability now ±0.3°C (was ±0.5°C)
- Klystron exchange (A2): Thales → EEV
- Various modifications at beamline endstations
- X10SA PX-II: Installation of Tell sample changing robot
- InnovAARE coordination with shutdown operations
- NEG coating area, preproject SLS renewal, evaluations with PSI divisions for SLS 1.5
- SLS 2.0 evaluations: machine, beamlines, area management, SECO visit
- Safety patrol (Sicherheitsrundgang) / new fire concept (evaluation)
- Implementation of PSI/GFA machine safety:
  - See https://intranet.psi.ch/de/safety-maschinen-roboter-sicherheit

---
Science highlight:
Structure of SARS-CoV-2 enzyme
- Measurements at the X06SA-PXI beamline by a team from the Goethe University in Frankfurt am Main, Germany, performed just before Easter (9 April 2020) reveal the structure and function of an essential enzyme of the SARS-CoV-2 virus («papaine-like protease»). The research was part of the COVID-19 priority call.
  - Media info: https://www.psi.ch/en/macromolecular-crystallography/scientific-highlights/first-mx-results-of-the-priority-covid-19-call
  - Publication: Shin, D., Mukherjee, R., Grewe, D. et al. Papain-like protease regulates SARS-CoV-2 viral spread and innate immunity. Nature (2020). https://doi.org/10.1038/s41586-020-2601-5

---
Machine science highlight:
booster chromaticity
(J. Kallestrup)
Nominal operation
- Recent studies at the SLS booster revealed a negative vertical chromaticity leading to a fast head-tail instability (bunch jitter and beam size blowup when increasing beam charge at smaller vertical beam size).
  - Simulation: (ξx, ξy) = (+5,+5)
  - Measured: (ξx, ξy) = (+7.9,–2.1)!
- The instability was present for 20 years!
- Now corrected by changing the sextupole family –4 A → –2 A
  - New measured: (ξx, ξy) = (+5.5,+4.8)
Improved operation

---
Wir schaffen Wissen – heute für morgen
Thank you for your attention