## Slide 1: SwissFEL Electron Diagnostics Review Workshop Oct 18 – 21, 2021
Reviewers’ Comments and Recommendations
Mario Ferianis, ELETTRA
Christopher Gerth, DESY
Patrick Krejcik, SLAC

## Slide 2: Generalities
The reviewers would like to thank the presenters for the excellent presentations
We present here our comments and recommendations with a slide for each of the talks.
We note that the requirements for the diagnostics depend on the modes of machine operation including charge, peak current, multi-bunch operation, interleaved bunches to multiple users etc. Future reviews would benefit from a brief introductory overview of operating modes and future plans.

## Slide 3: 1. Charge monitors – Gian Luca Orlandi
having redundant systems for charge measurement is a good move
overall ICT accuracy seems quite good (few pC in 200pC for most, with one exception)
  But accuracy measuring second bunch is a problem
  More bunches will require more changes
  It was not clear what the goal for accuracy was: 1% ?
Recommend synchronous acquisition between BPMs and ICTs for improved accuracy of comparison.
Cavity BPM seems to be “more reliable”, although not absolute
Recommend taking the sqrt of the energy in the BPM waveform as the charge signal for comparison e.g. rms of the waveform, as this is independent of any fitting routine.
Recommend periodic calibration check of the CM toroidal sensors, using calibration winding input, where available.

## Slide 4: 2. Electron BPMs – Boris Keil
BPM charge reading (monopole signal):
+ Accuracy better (in the 1% level) compared to the Turbo ICT (more on the 10% level)
+ No effect on bunch 2 as in the case of Turbo ICT
Recommend initially installing a few temperature controlled racks to test if this helps with drift issues
Relatively long (>5m) coaxial cables could be source to drifts vs temp/humidity
Charge measurements only available in dedicated measurements with oscilloscope
Recommend implementing charge readout in the processors (as in case of EuXFEL)
  SLAC uses the waveform rms method mentioned on the previous slide

## Slide 5: 3. Beam Loss Monitors – Gian Luca Orlandi
system is robust and well equipped
Nice MPS Graphical User Interface with lot of info and clear layout for immediate locking variable discovery
Fiber spools (DBLM), Cherenkov fibers (DLLM) and RadFET (DDRM) function well
But strongly recommend installing more of the fast fiber (DLLM) systems (Athos!)
Recommend against relying on RADFET for machine trajectory tuning (20sec/refresh)
Cherenkov fiber systems are typically much faster, almost real time, and allow for immediate losses identification with few meters space resolution
Improve spatial resolution of DLLM => new electronics with better ADC resolution (16 Bit instead of 12 Bit ?)
Readout of analog waveform, V(t), could provide better space resolution , like in OTDR
We learnt about strict Swiss limits on Max Dose / day
1Hz operation with screens is too slow for measurements and tuning
Recommend looking for creative ways to overcome this.
  E.g. 10 Hz burst mode?
  Shielding?

## Slide 6: 4. Porthos
This presentation clarified the diagnostics requirements
The overall machine layout is “quite complex”, with several “distributed” accelerating structures and undulators
Emphasized the need for multi-bunch discrimination in the linac
Especially when diagnosing jitter issues and you need to correlate downstream bunches with the correct bunch in the upstream linac
And also when measuring energy and compression for each of the multi-bunches
Having the different bunches generated by different laser oscillators in the Injector, adds extra relative jitter to the bunches
Using a single, high repetition rate, Laser oscillator and picking out individual pulses to the laser amplifier chains should mitigate this, with implicit benefit to the bunch stability
  Has an impact on BAM system

## Slide 7: 5. HERO – Pavle Juranic
Main diagnostic concern is the temporal overlap monitor
Acknowledge there budget constraints are of concern
If it eventually migrates from an experimental system to routine operations then we recommend some additional effort to replace the local oscilloscope control with integrated remote controls etc.
At Fermi, we routinely do coarse bunch to laser time alignment using few GHz photodiode, which is cheaper than PMT
Beside temporal drift monitoring, with holey screens, in “modulators” also the transverse (position & angle) monitoring is a key parameter , typically implemented with 2 screens (done at laser heater)
At Fermi we monitor the entry point with “virtual screen” and exit point with holey screen imaging (4 lobes)

## Slide 8: 6. Positron Production – P3
This system appears to be in the conceptual stage with some details still to be worked out.
The large dose anticipated for this experiment seems at odds with the strict safety limits (BAG) on dose for all other systems.
The Cerenkov fiber array may suffer from systematic errors as the incident angle changes and the critical angle for total internal reflection is exceeded in the fiber
Recommend a simpler, large area fluorescent screen to quantify the intensity distribution of the positrons. Moving the screen in z would give greater information on the phase space distribution of the positrons
The large aperture Faraday cup/beam dump is a non-trivial design
  2 fix FCs: one for e+ and one for e-
  Impedance matching in the design is taken care, however implementation in vacuum environment is non-trivial
  => we recommend a separate design review for the FCs before fabrication.

## Slide 9: 7. Synchrotron radiation monitors – Gian Luca Orlandi
The advantage of this monitor
  It is non-invasive
  For the gated camera: No ’cross-talk’ of bunches. Independent device to check ‘cross-talk’ of other devices: charge (intensity), compression (spread), beam energy (center-of-gravity)
But at other labs they are not used extensively
They provide some redundant information on longitudinal phase space from which bunch length can be calculated.
The linearity of the chirp, for example can be monitored from the intensity distribution.
Recommend improving the process data acquisition in control system
  to make the system more useful for routine operation and tuning (N.H.)
  to implement storing and archiving of processed data (center-of-gravity, spread, intensity) in EPICS
Recommend upgrading the GUI so that energy and energy spread are automatically calibrated and displayed on the plots in MeV
Recommend study correlation with DBCM (spread), DBAM (mean), ICT/DBPM (intensity)
Recommend 2nd DSRM station to be equipped with gated camera

## Slide 10: 8. Prof Monitor Screens – Pavle Juranic
(Problems with COTR/CODR from in-vacuum mirror solved)
Pleased to see the resolution issue has been solved with variable focus scheme
  Could also consider making all the filters have the same focal shift
  Re-do comparison of new DSCR with improved resolution and wire scanners
  (YAG does not seem to produce saturation effects – EuXFEL: LYSO replaced by GAGG:CE)
It is frustrating that motorized lenses don’t have easy control system interfaces
Rate limiting is a serious issue limiting the use of these systems
Recommend looking for thinner YAG screens of 20 or even 5 microns
  And please let us know!

## Slide 11: 9. Wire Scanners – Gian Luca Orlandi
Resolution and performance look good
Microfabrication of new wires looks very exciting for future low emittance beams
  We encourage further development
Our main recommendation is to work on 100 Hz operation for faster scans
  Dose rate is the same
And improve software controls for greater automation of wire scanner setup (automated PMT gain adjustment, scan ranges and speed, etc.)
  Wire scans should be a single button push to execute, and even operated by scripts
  Only then will Operators use them routinely and frequently for tuning and monitoring

## Slide 12: 10. Bunch arrival time monitors – Vladimir Arsov
Excellent resolution has been demonstrated with this system
Encourage you to complete the 2 bunch discrimination setup
It appears that the phase reference architecture is still in flux
Recommend that you plan now for all 4 future beam lines when you layout your phase reference distribution system, e.g. when putting in power splitters from the master laser oscillator, and design review
This class of BAMs definitely call for dedicated stabilized optical links.
Also, in a multiple laser oscillator photocathode arrangement, which is the ultra low phase noise “reference” ?
critical also to pump/probe laser oscillator synchronization

## Slide 13: 11. Transverse Deflecting Structures – Paolo Craievich
Excellent performance and resolution has been achieved with the S- and C-band systems
The PolariX TDS will be a valuable addition for FEL tuning and photon users
  Strongly recommend you pursue 100 Hz operation with the dump screen
  If this proves impossible an alternative is an off-axis screen where the TDS kicks and streaks the beam onto the screen at 1 or 10 Hz continuously
  LCLS photon users record every streaked image with their experiment data!
The passive streaker has impressive performance and clearly demonstrates the usefulness of a post-undulator streaker.
  A key advantage is the lack of jitter with self synchronization
  That said, it is still a “poor man’s” streaker compared to Polarix because of the nonlinear streak
We agree with your long term plans to install Polarix systems at the end of each undulator line
  A variable RF power splitter will enable you to share a single modulator/klystron to save cost
We agree that the passive streaker/dechirper would be more useful to keep upstream of the undulator where it can be used for both measurements and bunch manipulation in exotic setups like two-color operation

## Slide 14: 12. Bunch Compression Monitors – Gian Luca Olrlandi
Routine operation in conjunction with longitudinal feedback has been established
Two bunch discrimination has been solved with the MCT detector
The spectrometer is too slow to resolve multi bunches because of the multi channel detector
  Not needed for routine operation
  Recommend keeping this setup for future studies
  Spectrometer data may prove useful in future AI tuning of the machine

## Slide 15: Conclusions
This review demonstrates that PSI is a leading laboratory in the field of electron beam diagnostics and instrumentation, closely collaborating with RF & Machine Physics Teams
A very good collaboration level with main European labs also appears evident, in different areas (TDS, WS, BPM, Positrons etc)
Recurring themes in the presentations are:
  Limited repetition rate for measurements
  Improvements in automation of the measurements to allow routine and frequent use
  Adaption to multi-bunch discrimination with an eye to future beamlines
Phase reference architecture impacts diagnostics
Anticipate that AI and ML will play an increasing role in the future which increases the importance of non-invasive diagnostics even if they are used less frequently now (e.g. SRM, BCM spectrometer)
We look forward to continued collaboration and hearing about your progress!