## Slide 1: Beam size monitoring at SLS (2.0)
Cigdem Ozkan Loch :: Electron Beam Instrumentation :: Paul Scherrer Institute
23.03.2022

## Slide 2: Thank you!!
We have a habit in writing articles published in scientific journals to make the work as finished as possible, to cover up all the tracks, to not worry about the blind alleys or describe how you had the wrong idea first, and so on. So there isn’t any place to publish, in a dignified manner, what you actually did in order to get to do the work.
—Richard P. Feynman, Nobel lecture 1996
When art critics get together, they talk about style, composition, and technique.
When artists get together, all they talk about is where to get the best turpentine.
—Pablo Picasso

## Slide 3: Source Size measurement at SLS
Smallest beam measured with 𝜋-polarization (at 364nm) = 3.6±0.6um
Sources: several PSI presentations/reports

## Slide 4: π-polarization & interferometry @ SLS
The pi-pol simulation with SRW-IGOR and experimental results at SLS match very well
Both pi-pol measurement results are consistent with each other
Figures from: https://indico.cern.ch/event/306551/contributions/705064/attachments/583692/803462/12_SAA-ALERT2014.pdf

## Slide 5: Lessons learned at SLS
Only vertical beam size measurement
Radiation damage on vacuum window means measuring only once in half hour
When beam size measurement isn’t working everyone has a collective heart attack (yet, the machine still runs)
Easy to align
Beamline vibrations, due to cooling water, are amplified on mirror chamber >>> spatial oscillations and bad beam size readings
Mirrors do age in the UV light
Lens did not need replacing but mirror did
Vacuum window needs replacement once in 18 months
Have not been able to corroborate beam size so far >>> more on this later
Radiation will eventually dry out the grease of the linear stages which can cause loss of steps in case of alignment
Cameras age, servers die

## Slide 6: SLS Pinhole/π-pol chamber

## Slide 7: X-ray Pinhole
Electron beam
Dipole magnet
Diamond window
Air
Vacuum
d
D
Pinhole magnification: M = D / d
S = M · s
s
Ce:YAG
mirror
Camera & objective
Mo filters
Pinhole

## Slide 8: X-ray Pinhole: Point Spread Function
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20
sigma_y (um) | 27.9994 | 26.3003 | 25.5102 | 24.644 | 23.6731 | 23.4637 | 23.4827 | 23.7588 | 24.3156 | 25.3103 | 26.5335 | 27.8994 | 29.5938 | 31.3309 | 33.3442 | 35.2907 | 39.9264 | 42.2347 | 44.7571 | 47.1321
pi-pol | 11.8358 | 10.0756 | 9.6186 | 9.4166 | 9.5416 | 9.9883 | 9.9434 | 10.1761 | 10.2699 | 10.6982 | 12.0347 | 12.8668 | 13.6549 | 14.7265 | 14.6913 | 15.6644 | 18.1559 | 18.9815 | 19.894 | 21.5832
PSF of the pinholes for the 2.4 GeV (N. Milas, A. Saa Hernandez, 2013)

## Slide 9: Lessons learned at SLS
Pinhole is not well designed at SLS, not in use
Smallest beam size is always 20um (without de-convolving for pinhole size)
Only good for visuals: see beam tilt, beam blow up
Beamline vibrations not seen here as no components in between
Ozone is damaging: Have you ever broken a diamond vacuum window?
Optics turn brown
Cameras age

## Slide 10: SLS 2.0
BN VB
BE
Dispersion dominated
Emittance dominated
Nominal rms beam size (HxV) = 8.5um x 6.8um
Vertical emittance change of 1pmrad ~ 400 nm
Horizontal emittance of 7pmrad ~ 250nm
Horizontal emittance | Open IDs, low current | 150 pmrad
 | Closed IDs, 400mA | 136 pmrad
 | Nominal | 157 pmrad
Vertical emittance | 1-2% coupling | 2-3 pmrad
 | Nominal | 10 pmrad
 | Range of operation | 8-12 pmrad
Energy | 2.7 GeV
Vertical emittance | 10 pm.rad
Horizontal emittance | 157 pm.rad
Energy spread | 0.001145
My means are sane, my motive and my object mad.
—Captain Ahab, Moby Dick by Herman Melville

## Slide 11: SLS 2.0: Fresnel Zone Plates
Fresnel zone plates (FZPs) are used to build x-ray microscopes
Spatial resolution depends on the outermost-zone width of FZPs (100 nm is technologically achievable)
KEK-ATF damping ring uses FZPs in arrangement below to measure ultralow emittance electron-beam profile

## Slide 12: FZP: single vs. double
Electron beam
Dipole magnet
LC1
LC2
monochromator
Transmission ZP
Ce:YAG
mirror
microscope setup
CS
Electron beam
Dipole magnet
monochromator
Transmission ZP
Ce:YAG
mirror
microscope setup
CS
Magnifying ZP

## Slide 13: Central Stop and FZP Alignment
CS diameter = 1 mm
ZP diameter = 2.5 mm
ZP outermost zone width = 125 nm
ZP focal length = 2772 mm at 11keV
ZP height = ~ 2500 nm in Au

## Slide 14: FZP measurements at SLS
@ FZP focus, M = 0.205
µ-scope magnification = 20x
beam size = 22µm x 61µm
First image of ZP: simple placement into beam
- Zone plate clearly visible along with central stop
- Central high intensity spot is focused beam
Setup:
Monochromator set to 11keV
Source to FZP = 16.3m; microscope to FZP = 3.36cm
Zone plate f=2.77m, 2.5mm diameter, outermost width 125nm
µ-scope: Basler a2A-4504-18um PRO + 2x Mitutuyo lens + 5um Ce:YAG
In theory, theory and practice are the same. In practice, they’re different.
—Anonymous

## Slide 15
Chart: ZPX vs sigma_y
| -61 | -63 | -65 | -67 | -69 | -71 | -73 | -75
 | 52.74 | 51.74 | 50.5 | 49.14 | 48.15 | 48.57 | 49.38 | 50.55
Chart: ZPX vs sigma_x
| -61 | -63 | -65 | -67 | -69 | -71 | -73 | -75
 | 119.9 | 119.2 | 118.2 | 117.3 | 117.2 | 117.4 | 117.8 | 118.7
~30umx63um
Flipped YAG: ~23umx55um
Ask producer to mark the scintillating side

## Slide 16
Chart: 
| 2 | 1 | 0.6 | 0.2 | -0.2 | -1 | -2 | -3 | -4 | -5
sigma_y (um) | 32.0211 | 29.7889 | 29.7032 | 29.9031 | 29.6176 | 30.6504 | 33.1633 | 37.1993 | 41.5779 | 47.0274
Skew Quad current scan changes beam size
Pi-pol gives sigma_y ~ 9um
Scintillator doesn’t allow improvement
Use different substrate for 5um Ce:YAG? Glass?
How about AR coating?
>>> Trying a “wire scan”
Beam size on the scintillator ~6um

## Slide 17: FZP measurements at TOMCAT
min. horizontal beam size (rms) = 59.8um (optics 55.9um)
min. vertical beam size (rms) = 17.86um
TOMCAT beamline
@ 11.5keV, water cooled monochromator
Distances:
Source-FZP1: 17027mm
FZP1-det= 3510mm
Many thanks to my colleagues
A. Bonnin (TOMCAT)
N. Samadi (TOMCAT/X-ray Nano-optics)
J. Vila Comamala (X-ray Nano-optics)

## Slide 18: Effect of monochromator thermal deformation
Thermal deformation of the first crystal monochromator is a concern. If downstream beam focusing is planned, the lensing effect of the monochromator must be considered even if thermal deformations are small. The X-ray beam power results in an effective change of the focal distance; this is due to a lensing effect of the distorted monochromator that results in the creation of a virtual source downstream of the actual source.
On the influence of monochromator thermal deformations on X-ray focusing
https://doi.org/10.1016/j.nima.2016.02.103

## Slide 19: Thermal deformation ANSYS
Heat load 240 W
Many thanks to Dr. X. Wang

## Slide 20: Reduce heatload upstream
Heat load 7.5 W
Good, fast and cheap: pick any two.
—Anonymous

## Slide 21: 2 zone plate setup
X-ray source to Condenser FZP distance is 17 m
The 1st FZP de-magnifies the source about 5x.
Use the 2nd FZP to produce a magnified image. The FZP can have a diameter of 150 um and outermost zone width of 100 nm. Its focal length is 133 mm.
3313 mm
155 mm
932 mm

## Slide 22: Non-optimal TXM Setup
min. horizontal beam size (rms) = 75.1um
min. vertical beam size (rms) = 20.6um
Need larger Central Stop
>>> Do again with suitable optics
@ 11.5keV
Source-FZP1: 17027mm
FZP1-det= 3510mm
focus1-FZP2: 119mm
FZP2-Detector: 413.1mm
Image 1
Image 50
Image 21
Flat
Flat subtracted
Min x not reached
Difficult to subtract flat for the out of focus positions
A failed experiment can be more important than a trivial design.
—Verner Panton

## Slide 23: Beamline design
Kühlwasser Vorlauf
Kühlwasser Rücklauf
Pressure
FE-APAB1
Flow
FE-AB1WW
AB1
Flow
TC
FE-CSTC1
Pressure
FE-CSWP1
Pressure
FE-CSWP2
TC
FE-CSTC2
FE-CSVE1
FE-CSVE2
Druckluft
Mechanical
Vacuum
Water Cooling
Thermal
Pressurized Air
Flow
Flow
Flow
There is more madness to my method than method to my madness.
— Salvador Dali

## Slide 24: Monochromator
 | Actuator | Resolution | Reproducibility | Range
X | Remote controlled Move with chamber | 1um | 2um | 45mm
Y | Remote controlled Move with chamber | 5um | 20um | 50mm
Z | Remote controlled Move with chamber | 1um | 1um | ±5mm
Pitch | Remote controlled For wavelength/energy change | 0.1urad | 0.5urad | 8o to 12o
Roll | Not needed | NA | NA | 
Yaw | Not needed | NA | NA | 
1st Crystal
 | Actuator | Resolution | Reproducibility | Range
X | Move with chamber | 1um | 2um | 
Y | Move with chamber | 5um | 20um | 
Z | Move with chamber | 1um | 1um | 
Pitch | Move with 1st Crystal Fixed distance from 1st Crystal = 8mm | 0.1urad | 0.5urad | 8o to 12 o
Roll | Remote controlled For optimization | 0.1urad | 0.5urad | ±5o
Yaw | Not needed | NA | NA | 
2nd Crystal

## Slide 25: Monochromator
Mirror center of rotation
Beam height center (Z=0),
5mm behind the front edge
of lower mirror
X-ray beam
8mm x 56mm
Visible light
diameter = 56mm
1st crystal: 70x30x15
2nd crystal: 105x30x15
Distance between crystals fix 8mm
Page 25

## Slide 26: Summary
You can please some of the people all of the time, you can please all of the people some of the time, but you can't please all of the people all of the time.
—John Lydgates
Visible light for pi-polarization and bunch length measurements using streak camera
X-ray pinhole camera is easy to accommodate but PSF suggests not possible to measure small source size
Fresnel Zone plates for small source size measurements and for observing relative size changes
SLS tests with 2 zone plate design to be tested with suitable 2nd zone plate and central stop
Wire scanner to confirm the beam spot from the first zone plate
Does anyone know how to simulate thermal deformations and their effects in OASYS?

## Slide 27: But, that’s just what I think. What do you think??!?!? Thank you for your time, feedback and attention.