## Slide 1: Emittance Diagnostics at PETRA IV
Marco Marongiu
Hamburg, 2022.03.24
Elettra 2.0 Workshop on Optical Diagnostics for Low-Emittance Storage Rings

## Slide 2: Outline
01 PETRA IV Project
Main Characteristics
02 Emittance Diagnostics
Requirements
Beam Parameters
Radiation Sources
03 Measurement Techniques
Pinhole Camera (XPC)
Fresnel Diffractometry (XFD)
04 Conclusion and Future Developments

## Slide 3: PETRA IV Project
The PETRA IV project will be the successor of PETRA III, the 6 GeV synchrotron radiation facility in Hamburg.
It will be a Diffraction Limited Light Source with an emittance of 20/4 pm rad, instead of the 1300/10 pm rad currently obtained in PETRA III.
The new machine will operate at the same beam energy (6 GeV) of PETRA III with an intensity of 200 mA (twice the value of PETRA III).
The machine will be installed in the same tunnel of the current operating one, but a new experimental hall will be build increasing the number of available beamlines.

## Slide 4: PETRA IV Project

## Slide 5: Emittance Diagnostics Requirements
Two X-Ray beamlines for measuring emittance and energy spread:
  Low Dispersion contribution
  High Dispersion contribution
Emittance Resolution of 0.5 pm  Beam size resolution of 100 nm
High Dynamic Range:
  A measurement for beam size in the range1-18 µm (i.e. Fresnel Diffractometry: XFD)
  A measurement for beam size above 15 µm (i.e. Pinhole Camera: XPC)
Screen outside the accelerator tunnel  Beamline length around 45-50 m

## Slide 6: Beam Parameters
ID | L [m] | B [T]
DL | 0.303 | 0.28
DQ | 1.84 | 0.19
Emit_x=20pm; Emit_y=4pm;
E_spread=1e-3
ID | S [m] | BETX [m] | BETY [m] | DX [m] | spot_x [µm] | spot_y [µm] | Dispersion Contribution [%]
DL | 1096.287 | 2.07 | 8.57 | -0.0001 | 6.4 | 5.8 | 0.01
DQ | 1090.6 | 2.19 | 7.96 | 0.02 | 23.9 | 5.6 | 72.3
Difficulties of extraction of the photon beam from DQ
  Canting solution under study:
    3 Pole Wiggler as a radiation source

## Slide 7: 3 Pole Wiggler
Magnetic gap 22mm
Peak field: 0.86 T
PM material: Sm2C017
Magnetic length: 108mm
Courtesy of G. Le Bec (ESRF)
Taken from J. Chavanne, DLSR2016
Compared to BM:
Higher intensities
Same resolution

## Slide 8: X-Ray Beamlines Sketch
Overall length of ~50 m  Screen outside the accelerator tunnel
Pinhole (XPC) for the high dynamic range (beam above 15 µm)
Diffractometry (XFD) for small beam (1-18 µm)
  Monochromator (i.e. Si(311)) is required for the XFD measurement
X-Ray Source
XPC
(d~1.5 mm; A~20 µm)
XFD
(d~1.5 mm; A~80 µm)
Monochromator
(~16 keV)
Scintillator
Screen
~10 m
~20 m
~20 m

## Slide 9: XFD Theory
If A is optimized after each variation of wavelength, L or R:
Ratio is constant (~0.38).

## Slide 10: XFD PSF
Reoptimize the aperture for each value change.
M=1
M=2
M=3
E=8 keV
E=12 keV
E=16 keV

## Slide 11: XFD Resolution vs photon energy
Different Pinhole Position
M = 1
Reoptimize the aperture for each E value.
Peak to Valley Ratio = S; #bit = N
10 bit
σ=8 µm

## Slide 12: XFD Resolution vs magnification
Different Pinhole Position
E = 16 keV
Reoptimize the aperture for each M value.
Peak to Valley Ratio = S; #bit = N
10 bit
σ=8 µm

## Slide 13: XFD Resolution (10 bit CCD) L=25m
Different Photon Energy
Different Magnification
Reoptimize the aperture for each Energy and Magnification value.
Minimum beam size resolvable (µm)
Maximum beam size resolvable (µm)

## Slide 14: XFD Dynamic Range(10 bit CCD) L=25m
Different Photon Energy
Different Magnification
Reoptimize the aperture for each Energy and Magnification value.
Minimum beam size resolvable (µm)
Dynamic range(µm)

## Slide 15: XFD Dynamic Range (10 bit CCD) L=30m
Different Photon Energy
Different Magnification
Reoptimize the aperture for each Energy and Magnification value.
Minimum beam size resolvable (µm)
Dynamic range (µm)

## Slide 16: XFD: Beamline length of 50m
8 keV Case |  |  |  |  | 
L (m) | M | Min Spot (µm) | Max Spot (µm) | Peak Distance (µm) | Aperture size (µm)
25 | 1 | 3 | 30 | 67 | 116
30 | 0.6 | 2 | 25.5 | 55 | 114
35 | 0.4 | 1.3 | 21.1 | 43 | 107
40 | 0.25 | 0.8 | 16.6 | 25 | 93
12 keV Case |  |  |  |  | 
L (m) | M | Min Spot (µm) | Max Spot (µm) | Peak Distance (µm) | Aperture size (µm)
25 | 1 | 2 | 25.4 | 55 | 95
30 | 0.6 | 1.4 | 21.4 | 45 | 93
35 | 0.4 | 0.9 | 17.6 | 37 | 87
40 | 0.25 | 0.5 | 13.8 | 23 | 76
16 keV Case |  |  |  |  | 
L (m) | M | Min Spot (µm) | Max Spot (µm) | Peak Distance (µm) | Aperture size (µm)
25 | 1 | 1.5 | 22.6 | 48 | 82
30 | 0.6 | 1 | 18.9 | 40 | 81
35 | 0.4 | 0.7 | 15.4 | 32 | 75
40 | 0.25 | 0.4 | 12 | 22 | 66
L=30; M=0.6; E=16 keV:
min=1, max=18.9, peak distance=40, A=81 µm

## Slide 17: XPC: D=50m
L=10 m; R=40m; E=16 keV  A=24µm; PSF=11µm
L=25 m; R=25m; E=16 keV  A=30µm; PSF=24µm
Higher photon energies and higher magnifications means better resolution.
But the aperture becomes small and too difficult to manufacture.

## Slide 18: Pinhole Manufacturing
High aspect ratio (thickness > 1mm; aperture~25µm):
LIGA techniques (DIAMOND) thickness~200nm
Laser Drilling  Prototype testing foreseen
Different Material to be tested:
Tungsten (i.e. PETRA III)
Tungsten Carbide (i.e. ESRF)
Gold (DIAMOND)
Experimental PSF evaluation (i.e. in Touschek dominated beams the vertical size is proportional to the beam lifetime)
Condition Number (h-1 mA-1)

## Slide 19: Conclusion
XFD for high resolution measurement and XPC for high dynamic range
Overall length of ~50 m  Screen outside the accelerator tunnel
Bending magnets and 3PW are under consideration as radiation source
Future Developments
Pinhole characterization:
  Quality of the pinhole (aspect ratio, materials, manufacturing techniques)
  Experimental PSF evaluation:
    Further test at PETRA III
XFD beam test
Optical Interferometer
Streak Camera for longitudinal measurements

## Slide 20: Thank you

## Slide 21
Marco Marongiu
MDI
marco.marongiu@desy.de