SwissFEL Porthos, visions for the future
Thomas Schietinger :: Paul Scherrer Institut
Machine report week 13/2023
SwissFEL Exchange Meeting, 20 March 2023

---
Calendar

---
Photon (pulse) energies
 ● Photon energy (keV)
 ● Photon pulse energy (µJ)
Mon,13-Mar, 08:00 – Mon, 20-Mar, 08:00
Aramis
Setup
9.5 keV for Bernina, ~1.1 mJ
9.5 keV for Bernina, ~1.0 mJ
6.5 keV
for Alvra,
~1.0 mJ
12 keV
for Alvra,
~480 µJ
reference
measurement
at 500 eV
standard beam
without tilt and with
varying beam tilts
Athos
back to multistage
amplification mode
(710±50 eV)
Setup
710 eV for Maloja in multistage
amplification mode, ~650 µJ

---
E. Prat, T. Weilbach,
Ch. Kittel, N. Hiller
Aramis setup
- Setup for 9.5 keV starting from a reference file taken at 9.3 keV.
- Initial optimization gives 1.2 mJ.
- After some work on the gun laser still 1.1 mJ.
- Spectrum not perfect (shoulder at higher energies)
  but acceptable for the users.
  - Focus on Athos setup.
  - Special mode in Athos leaves very few knobs
    to improve the spectrum after initial setup...
Bernina monochromator scan:

---
E. Prat, G. Wang,
T. Weilbach, Ch. Kittel
Athos setup
Monday/Tuesday:
- Set up fresh-slice multistage amplification at target energy of 710 eV with four stages
  (7+3+3+3 undulators). About 600 µJ in 2–3 fs (FWHM).
- Check energy scaling up to ±50 eV.
Friday:
- Setup of standard beam at 710 eV (no tilt, 40–50 fs FWHM, 3.8 mJ), and three pulse lengths based on beam tilt:
  - small tilt:	20–25 fs FWHM (1.56 mJ)
  - moderate tilt: 10 fs FWHM (650 µJ)
  - large tilt:	8 fs FWHM (235 µJ)
Sunday:
- Reference measurement at 500 eV (full beam)
- Return to multistage amplification at 710 eV (energy scaling works less well than earlier)

---
E. Prat, G. Wang,
T. Weilbach, Ch. Kittel
Athos setup
4 stages:
Fresh-slice multistage amplification at target energy of 710 eV with four stages (7+3+3+3 undulators)
3 stages:
1 stage:
2 stages:
No lasing:
time
energy

---
Issues
During the whole week:
- Some instability (charge jitter) in Aramis due to instability in gun laser (Alcor) since the exchange of a control unit last week. Was addressed this morning.
- Instability from BC2 compression monitor (CDR): state machine does not retract filter automatically after a beam hickup → requires manual intervention by operators
- Some problems with BC1 feedback for bunch-2 (Athos) (actuator jumps, to be investigated…). Was left off during multistage operation.
- CHIC settings not restoring properly after a change in K value – under investigation (D. Voulot)
- Frequent trips from SATCB01 RF station (Athos linac). Needs more conditioning. For now accelerating voltage reduced from 230 MV to 220 MV, this improved the situation a lot.
- Enhanced beam losses in Athos when operating with tilted beam. BLM thresholds temporarily raised, DRMs within allowed range.
- Significant beam losses in Athos with multistage amplification. Allow only 50 Hz during measurements, 10 Hz between measurements.
Specific events:
- Thursday, 17:00: SINSB02 RF station down, replacement of a relay.
- Sunday, 22:00: SATUN22 undulator gets stuck, needs a reset.
- Monday, 3:30: failure of a turbopump at the Athos gas detector, no beam in Athos (Maloja users had already left, no vacuum pikett for photon components, therefore just wait...)