## Slide 1: Magnetic measurement activities at CERN Part I: projects, infrastructure and collaborations Part II: instrumentation and techniques
Marco Buzio on behalf of the Test & Measurement Section
International Magnetic Measurement Workshop IMMW23, Bad Zurzach, 07.10.2023

## Slide 2: Projects, infrastructure and collaborations
Ongoing and upcoming

## Slide 3: High-Luminosity LHC
Highest-priority CERN project:
replace rad-damaged IR quads
10 × luminosity → +25% discovery range
LS3 (shutdown) schedule under review
Model & prototype tests, Horizontal test benches upgrade
Cryomagnet series tests String 2 tests

## Slide 4: HL-LHC magnet tests
~50 cryoassemblies, 14 types of main/corrector magnets
First operational accelerator-quality Nb3Sn magnets
Magnetic measurement challenges: - Up to 109 mm anticryostat bores, 15 m rotating coil systems - high measurement accuracy required (IR region): 5 units absolute B2, 0.1 units multipoles - 3 warm rotating coil systems to be maintained remotely on the manufacturer’s premises (2 units B2, 0.1 units multipoles) - warm/cold magnetic vs mechanical axis, longitudinal alignment of IR triplets (~5 mm)
23 Nb3Sn triplet quads (MQXF) 150 mm bore, 4.2/7.15 m, 132 T/m (~11 T peak) @ 16.2 kA
7 NbTi separation/recombination dipole (MBXF ) 136.7 mm, 6.3 m, 5.6 T @ 12 kA
Mariano Pentella this IMMW

## Slide 5: SM18 cryogenic test station upgrade
4/10 horizontal benches upgraded for HL-LHC cryomagnets, 1/10 for SC Link, 1/10 for coil shaft storage
Preparations for String 2 under way (25 → 60 g/s Lhe liquefaction, 14 g/s pumping @ 12 mbar)
Power circuits upgrade: 2×20 kA, 10V converters, 4×2kA auxiliary circuits, load and polarity switches
DAQ and protection: quench heathers, energy extraction switches, UQDS, CLIQ, PLC & fast interlock systems, anticryostat heaters …
SC Link/DFX on Bench F2
MQXFB04 on Bench F1
Patch Panel Interface/DFHX

## Slide 6: FCC - Future Circular Collider
Z, W, H factory with highest luminosity, lowest cost and risk for E frontier collider
feasibility study ongoing, momentum rapidly building up
key roadmap milestones: 2026: EU strategy update ~2028: approval ~2045: FCC-ee commissioning 2070, FCC-hh operation
L=90 km, ρ=10.7 km
e-e collider: E=46 to 182 GeV, 2900 24 m bending dipoles with B1=14 .. 57 mT, 2900 3.1 m, 10 T/m focusing quads
e-e booster: B1=6.3 mT (inj.) ... 58 mT (ext.)
Prototype e-e collider bending dipole

## Slide 7: Prototypes for FCC
Credit: Carlo Petrone
SuShi (passive MgB2 shield for 3T FCC-hh septa)
FCC-ee magnets
10  8 Arepoc HHP-NP (200 mV/T @ 20 mA) cross-calibrated at RT (relative measurement)
8.3 mm MgB2 thickness after reactive Mg infiltration
flux jumps at lower field on ramp-down
1270 mm
gradual penetration from end

## Slide 8: HFM – High Field Magnets program
Recent project reorganization, including clear setting as direct R&D towards FCC-hh
Collaboration with PSI, INFN, Uppsala, CIEMAT, CEA and several universities
Program highlights:
Robust Cos Nb3Sn dipole up to 12 T (FalconD);
Target: 14 T dipole for 90 TeV FCC-hh in ~2050
Roadmap for feasibility of HTS for accelerator magnets
Non-FCC activities: 20/40 T Muon Collider solenoids
Vincenzo Di Capua this IMMW
Tests of electronic sensors and acquisition chain components in LHe @ 1.9 K and 4.5 K
Notes:   Selection, test, and calibration of different electronic components and sensors for use at cryogenic temperature
  First results are available (1st campaign) for
    MEMS microphones
    Piezoelectric transducers
    Amplifiers
    Analog multiplexer
  Other components will be tested (2nd and 3rd campaign in 2025):
    Fiber optic transmitters and receivers
    Integrated circuits (FPGA, ADC)

## Slide 9: HFM: Racetrack Model Magnet (RMM)
Chart: Central Field in T
RMM 1a | 15.259 | 15.935 | 15.987 | 16.044 | 16.316 | 16.304 | 16.383 | 16.441 | 16.392 | 16.411 | 16.423 | 16.441 | 16.461
RMM 1b | 15.968 | 16.183 | 16.236 | 16.28 | 16.346 | 16.4 | 16.439 | 16.498 | 16.496 | 16.201 | 16.206 | 16.361 | 16.346
RMM 1c cycle 1 | 14.682 | 15.674 | 15.722 | 16.06 | 15.953 | 16.329 | 16.548 | 16.196 | 16.154 | 16.179 | 16.525 | 16.286 | 16.185 | 16.215 | 16.388
RMM 1c cycle 2 | 15.681 | 16.072 | 16.315
Family of demonstrator magnets based on eRMC (racetrack coils) to explore full potential of Nb3Sn
RMM1d designed to reach a peak field of 16.7 T, in a 50 mm aperture, at 12 kA
Credit: Carlo Petrone, Mariano Pentella, Melvin Liebsch

## Slide 10: Physics Beyond Collider Experiments
SHiP
Search for Hidden Particles (dark matter/neutrinos) detector based on CERN SPS beam dump
Needs a 0.15 T, 4×6×3.5 m2 spectrometer that would consume 1.2 MW if resistive → MgB2 option
AION
Atom Interferometer Observatory and Network for dark matter/gravitational wave detection in PX46 pit
B field fluctuations → systematic errors
In-situ measurements DC-300 Hz with a fluxgate

## Slide 11: Other Superconducting Magnet Projects
EESD Spectrometer demonstrator
20K, MgB2 Energy Efficient Superferric Demonstrator (for SHiP experiment) @ 4000 A DC without ohmic losses
Phase I/II tests with cold coil/yoke at 1.9 to 30 K completed; Phase III tests with warm yoke scheduled soon
Field dynamics measured with standard rotating coils + hybrid coils/Hall sensors embedded in the coils
FUSILLO demonstrator
Magnet demonstrators of strongly curved CCT NbTi superconducting dipole for medical application gantries
Two subscale models with 2-turns coil fully tested. Full-scale 3-T model under construction and scheduled for testing by end of 2024.
Credit: Carlo Petrone, Mariano Pentella

## Slide 12: Hadron therapy gantry: GaToroid
12
Concept: toroidal gantry for hadron therapy, replacing rotating components with 1/r tokamak-like toroidal field
Initial models tested: - ReBCO version, powered only up to 15 A at 20 K - NbTi version, powered above nominal without performance limitations
Measurements performed at 1.9 K and 4.5 K via hybrid sensors consisting of PCB coil and Hall sensor
Credit: Luca Bottura, Carlo Petrone, Mariano Pentella

## Slide 13: Refurbishment/upgrade of resistive magnets
Ageing stock of 5000+ beam line normal-conducting magnets (EOL bathtub curve)
Common failures: water leaks, insulation integrity → consolidation and preventive maintenance
Occasion for power consumption optimization: DC → pulsed excitation
SPS fixed-target East Area 58 magnets in 2021
SPS fixed-target North Area (planning under discussion)
new PS extraction bumper with pole face loops for passive sextupole compensation
LIU – 120 new/refurbished magnets tested in 2016-2020

## Slide 14: Collaborations
Real-time magnetic measurements hysteresis modelling
Vertical anticryostat for Bench Cluster D up to 600 mm × 5.5 m magnet
1 × 30 kA + 2 × 2kA converters
Pawel Kosek this IMMW

## Slide 15: Instrumentation highlights
Sensors, probes and methods

## Slide 16: Rotating coil systems for HL-LHC
  ROHACELL filling (RESARM)
  PCB coil array
  Roller with ceramic ball-bearings, Ti bellows
  C fiber shell
  Nano connectors
Guy Deferne this IMMW
Mariano Pentella this IMMW
Credit: Lucio Fiscarelli
  Quadrupole calibration setup based on CERN-standard rotating coil system
  Mole for warm measurements

## Slide 17: Modelling and workflow tools
Python-based software framework to achieve fully traceable measurements from calibration, data acquisition (FFMM C++ framework) and analysis to the final measurement report.
Combine FEM/BEM simulations (ROXIE, Ansys, Comsol …), analytical models and measurements to update magnet models to an as-built state in a DB
Matthias Bonora this IMMW

## Slide 18: Flexible-PCB Quench Antenna
Novel design concept: rolled-up flexi PCB, very easy to (dis-)assemble
Optimized for MQXF (hard-wired analog bucking)
Improved G=100~500 acquisition electronics (analog-bucking op-amp bypassed, Zin adapted to long coax capacitive load)
B4
A4
B3
A3
2 × double-layer flexi PCB glued together 4 × 16-turn tangential coils per layer
12 ×Ø100 mm, 0.6 m long segments
Per-layer harmonic bucking
Credit: Lucio Fiscarelli

## Slide 19: Quench Antenna with twisted coils
B3, A3 B4, A4 ,
On one segment there are 3 sets of coils:
- Straight coils sensitive to B3,A3 and B4,A4 - Additional twisted coils sensitive to B3,A3
Flux jump/quench events modelled as elementary magnetic moments
Transversal position obtained from B3,A3 and B4,A4 measured with straight coils.
The longitudinal position z is proportional to the phase shift of the signals from twisted coils respect to the signals from the straight coils.
= twist pitch
Credit: Lucio Fiscarelli

## Slide 20: Quench Antenna results
Heatmap of flux jumps during an energization ramp (darker color indicates more events)
  Low field: large number of jumps distributed mainly in the inner layer
  At intermediate level, they tend towards the midplane of outer layer
  From intermediate to high fields, the number of jumps decrease
Credit: Lucio Fiscarelli

## Slide 21: Translating fluxmeter for solenoid mapping
PCB-based translating fluxmeter to map solenoids (motivated by CERN AD electron cooler)
New design optimized for 60 mT, includes 20 azimuthal local coils/ring for improved axis and field error measurement.
Fiducialization still a challenge. New approach based on 3D cylindrical harmonic description is being developed.
Credit: Mariano Pentella Carlo Petrone

## Slide 22: Strongly curved PCB fluxmeter (FUSILLO)
Piecewise PCB size < 500 mm for design flexibility, much lower cost and faster procurement
Dense point-like local field probes for cross-calibration with induction coils, missing gap compensation in non-uniform field
More complex mechanics, assembly and wiring procedure
Credit: Mariano Pentella, Unai Martinez

## Slide 23: 3D Hall probes
ASENSOR HE444 3D Hall sensor (the last one !?) 8 mrad orthogonality, 0.3% non-linearity
Piezo motor-based Spherical harmonic calibrator
CERN/NIKHEF 3D Hall Transducer (courtesy N. Pacifico, EP-DT)
Permanent-magnet double-cone fiducialization reference
Credit: Melvin Liebsch

## Slide 24: Sensor fusion: induction coils + Hall probes
Trend: combine both sensors on same PCB, whenever possible
Trivial applications: magnetic length, center field, fringe profiles
Real interest: cross-calibrate by exploiting coil’s linearity, Hall sensor’s absence of drift
Nb3Sn racetrack model coils
Detection of dynamic effects
3× aerospace-grade LakeShore HGT 2101 Hall probes (highly nonlinear!)
Fast ramp-down→ Hall sensor calibration
coil drift 100 ppm/s
raw flux, corrected
Credit: Lucio Fiscarelli, Carlo Petrone

## Slide 25: Non-magnetic material properties
Flux distortion method for very low permeability material (low µr → high field) @ room temperature
Analytical treatment possible for simple geometries; arbitrary samples need FE simulations
Typical accuracy 100 ppm, repeatability 10 ppm (best result: µr = 1.00085 of a W alloy sample, validated by vibrating sample)
1 T dipole, high uniformity background field
Translation stage
Moving NMR probe (5 ppm accuracy)
Sample: Ti hip prosthesis
(FE model output)
(measured)
(measured)
Worst-case (Fe-Cr prosthesis): µr=1.0023 Fm≈mg for BB≈42 T/m2 (i.e. ~ 20 T MRI magnet!)
Historical prosthesis samples
Credit: Mariano Pentella

## Slide 26: FEM/BEM processing of boundary data
Btangent
Bnormal
Back to Maxwell: Unique solution of Laplace equation in a source-free domain + Neumann (normal field) or Dirichlet (tangential) boundary conditions (if: simply connected domain, smooth boundary)
Reduce problem dimensionality: scan only the boundary, fit a BEM model to get interior values at arbitrary resolution (no volume mesh required !)
Back to Maxwell: Maximum Principle: extremal values always on the boundary → metrological advantage of interpolating the interior
R&D challenges: automated meshing and setup of BEM model; hybrid boundary conditions (e.g. integrals from flux measurements); incorporate arbitrarily scattered/multiple/interior points
BEM reconstruction
Full grid
2D example, cylindrical symmetry: Measurement uncertainty at |z|=r0
Credit: Carlo Petrone, Melvin Liebsch

## Slide 27: Hybrid approach: data-driven Digital Twins
Combine optimally the strong points of computer modeling and measurement
Formalize and automate the comparison of computed and measured quantities
Example: CERN ISOLDE 90° High Resolution Separator
beam pipe
Sensors in the fringe field region
computation
Parameters p
(g, µr)
Bayesian priors
(g0,σg), (µr0,σµr)
… and we want the field here …
… but we can only measure Bj there !
Measurement uncertainty
σB
measurement
Minimization w.r.t. p
Indications for experiment design
We don’t know well these …
Melvin Liebsch CAS on Magnets 2023

## Slide 28: Hysteresis measurement and modelling
Instrumentation R&D to measure integral dynamic effects in fast-pulsed, iron-dominated magnets (CERN PS, SPS)
EPA (Efficient Particle Accelerators) initiative: exploit Machine Learning to feed-forward hysteresis effect to machine control
LLMs trained on B-train measurements provide excellent results on cycles within training parameter space envelope
Vincenzo DI Capua this IMMW
Credit: Vincenzo Di Capua, Mariano Pentella, Anton Lu
Integral coil setups for the main PS/SPS magnets
Discrete PCB array + local coils/Hall probes
Just two copper bars !
Feed-forward correction of SPS operation

## Slide 29: Conclusions
Spectrum of existing instrumentation technologies covers well all current requirements, with widely variable levels of accuracy, difficulty and cost
Future demands are expected to push the envelope in terms of field strength, ramp rates, magnet size …
Waiting for the definition of a future strategy, we must prepare to test magnets over a vast range of parameters