## Slide 1: Latest Developments in Magnetic Measurements
Julien Songeon
IMMW - October 8th, 2024
Notes: Hello everyone, I’m Julien physicist from Metrolab and today I’ll present our [title]

## Slide 2: Outline
1. About us
  Small introduction
2. NMR Measurements
  Continuous wave
  Frequency modulation
  Frequency sweep
  Pulsed wave
3. Magnetic Field Camera
  Shimming
  Motorized MFC 8246
0. Outline
Notes: For this presentation,
I’ll briefly introduce Metrolab and the main instruments that we produce
I’ll start by presenting the improvement on our NMR method by moving from continuous wave to pulsed wave measurements.
I’ll follow up with improvement that we made on a motion stage mapper for shimming small bore magnet

## Slide 3: Metrolab Technology at a Glance
Founded in 1985 in Geneva - from CERN to MRI
Global market leader for high precision magnetometers
Provide equipment to medical imaging, research laboratories, and industrial manufacturing
High precision magnetic field measurement instruments
  NMR Teslameter & Camera
  3D Hall Magnetometer
  Fluxmeter (Digital Integrator)
ISO 17025 accreditation for NMR
1. About us

## Slide 4: Metrolab Technology Main Products
Chart: 
NMR Magnetic Field Camera (MFC)
NMR Precision Teslameter (PT)
Calibration & Service
Components
Three-axis Hall Magnetometer (THM)
Fast Digital Integrator (FDI)
1. About us
3.
2.
Notes: Our main products are[list them]
And today I’ll present the novelties in the left column

## Slide 5
Continuous Wave and Pulsed Wave
Notes: Our former precision Teslameter PT 2025 continuous wave, PT2026 pulsed wave

## Slide 6: PT2025: Continuous Wave
~
V
Detection
Peak
V
20 MHz
Counter
1s
t
Pulse if:
Integrator
2. NMR measurement
PT2025 is an analogical instrument
Measure the magnetic field using NMR with continuous wave excitation
Notes: Here is a general diagram that describe how the NMR measurements are performed.
First, lets start with the NMR excitation part

## Slide 7: PT2025: Continuous Wave
The excitation of the sample is performed by generating a fixed frequency continuous wave
  If excited at the Larmor frequency: signal absorption
A modulation of the field is applied on top of the excitation by a second coil
  By the superposition principle, the total magnetic field shifts with the modulation
2. NMR measurement
Notes: stimulating continuously at the NMR frequency a coil that surround the NMR sample.
By creating a local variation of the magnetic field using a modulation coil.

## Slide 8: PT2025: Continuous Wave
If the excitation frequency is not at the Larmor frequency, the sample will not be excited and the signal will not change.
If the frequency is at the Larmor frequency, the signal will be absorbed by the sample.
2. NMR measurement
Notes: slowly adjust the frequency until we “tune in” to resonance.
we must be certain that the NMR frequency is related to the magnetic field that is measured and not a shifted version of it.

## Slide 9: PT2025: Continuous Wave
The excitation frequency is lower than the Larmor frequency
The excitation frequency is higher than the Larmor frequency
We want to have the resonance when the modulation is null, such that the unperturbed B0 is measured
The excitation frequency is at the Larmor frequency
2. NMR measurement
Notes: This problem is solved by making sure that the NMR absorption peak occurs when the current injected in the modulation coil is null.
If the CW is not exactly at the Larmor frequency, the resonance will happen when the modulation is not null.
To correct this, [next slide]

## Slide 10
Analysis of between the absorption and the zero crossing of the triangular modulation wave
: NMR lock
t
Detection
peak
V
Pulse if:
2. NMR measurement
PT2025: Continuous Wave
Notes: A pulse is emitted at the moment the signal is absorbed and another one when the current is null.
Knowing the delta time between the two event, and knowing the speed of the modulation, it is possible to know the Larmor frequency

## Slide 11: PT2025: Continuous Wave
1Hz measurement by counting the frequency over 1s.
Analog feed-back loop regulation
~
V
Detection
Peak
V
20 MHz
Counter
1s
t
Pulse if:
Integrator
2. NMR measurement
Notes: The feedback loop will adapt the excitation frequency to follow the magnetic field

## Slide 12: Magnetic Field Camera: MFC 3045
Camera: array of NMR probes
MFC 3045 is also analogic and based on continuous wave
Used to map magnetic field to obtain the DSV
2. NMR measurement
Notes: As introduced, the MFC is an array of NMR probes

## Slide 13: MFC 3045: Continuous Wave
Uses a sweep of frequency modulation directly in the excitation coil
No modulation coil
Multiple NMR probes
OCXO
DDS
Detection
Peak
Detection
Peak
Detection
Peak
?
t
Counters
2. NMR measurement
Notes: This instrument no longer have a modulation coil and uses directly a modulation of the excitation frequency through its excitation coil
Let’s look again at the NMR part of the diagram

## Slide 14: MFC 3045: Continuous Wave
The Direct Digital Synthesis (DDS) produces the variation of the excitation frequency from to with a given increment.
DDS
Detection
Peak
2. NMR measurement
Notes: The DDS [read slide]. The sweep will be performed once rising and once falling

## Slide 15: MFC 3045: Continuous Wave
Detection threshold
2. NMR measurement
Notes: If the Larmor frequency is within the sweep frequency range, the sample will absorb the excitation signal and emit a signal in return.
Using a detection threshold, one can detect when this occurs.
We obtain a time t1 which is the time between the start of the rise and the moment of signal detection, and a time t2 which is between the start of the fall and the second detection

## Slide 16: MFC 3045: Continuous Wave
The excitation signal is shifted downward relative to the Larmor frequency
The excitation signal is shifted upward relative to the Larmor frequency
2. NMR measurement
Notes: If […] the first excitation will happen after a longer period of time, and the second period a shorter one. And if it is […] it will be the opposite the first time will be shorter and the second time longer.

## Slide 17: MFC 3045: Continuous Wave
Knowing the value of the increment, we can find the two frequency with:
The average of the two frequencies gives us the frequency measured by the MFC 3045
2. NMR measurement
Notes: The averaging between the rising and the falling measurement makes it possible to limit the errors linked to the voltage threshold for detection.

## Slide 18: MFC 3045: Continuous Wave
The measure is then performed on each probe
Unlike the PT 2025, the MFC 3045 does not disturb the field to make the measurement.
MFC3045 uses open-loop whereas PT2025 uses feedback loop
OCXO
DDS
Detection
Peak
Detection
Peak
Detection
Peak
?
t
2. NMR measurement
Counters
Notes: Feedback loop: constantly correct the excitation frequency, open-loop will go to the next set of frequency interval if the excitation did not happened

## Slide 19: PT2026: Pulsed Wave
PT2026 is a digital instrument that uses pulsed wave
The same coil is used to send the excitation signal and to received the signal from the sample
Has a precision down to 10ppb (parts per billion)
2. NMR measurement
Notes: a broad-band pulse, and the sample absorbs and re-emits at the Larmor frequency

## Slide 20: PT2026: Pulsed Wave
The method is based on the Free Induction Decay (FID) where the signal generated by an excitation pulse is followed by the acquisition of the signal decay.
If the excitation frequency is at the Larmor frequency, the signal will oscillate and will decay exponentially according to T2*
2. NMR measurement
Notes: a broad-band pulse, and the sample absorbs and re-emits at the Larmor frequency

## Slide 21: PT2026: Pulsed Wave
The Fourier transform of an oscillating signal with exponential decay is a Lorentzian function
2. NMR measurement
is the exponential decay constant (
is the Larmor frequency
Notes: the exponential decay will influence the width of the frequency peak
The maximal point of the peak will be the Larmor frequency

## Slide 22: Continuous and Pulsed Wave: Summary
Moved from continuous to pulsed wave
  requires fast-switching modern electronics but is more direct and offers greater precision and information.
  Faster measurements from 1Hz to 33Hz
  No field perturbation with modulation coil
  Improvement of the resolution by a factor 7.3
Digitalization
  Numerical computations
  Fourier transform
  Frequency analysis
2. NMR measurement

## Slide 23
Magnetic Field Camera
MFC 8246
Notes: This our new Magnetic Field Camera: the 8246
Used to for the shimming of small bore magnet. What is shimming and what is it needed?

## Slide 24: Shimming a magnet
3. Magnetic Field Camera
Shimming is the process to make B0 field of a magnet more homogeneous
Two types of shimming
  Passive shimming: small pieces of metal are placed within the bore of the magnet
  Active shimming: compensation of inhomogeneities with shim coils
W. Wang et al., 2023
doi:10.1002/mp.16538
Notes: Shimming is the process of homogenizing a magnet
There are two types of shimming: [read slide]
Distorted the field

## Slide 25: Shimming effect
3. Magnetic Field Camera
Y. Terada et al., 2013
doi:10.1016/j.jmr.2013.02.005.
Effect of passive and active shimming on MRI images
  Without shimming, the images are not usable
  With passive shimming only, images become useful and only small inhomogeneous areas remain due to susceptibility effects
  With both active and passive shimming, image artifacts are sufficiently suppressed so that anatomical structures are clearly visualized
Passive shimming is needed as active shimming alone is limited in the compensation of large field inhomogeneities
Notes: Shimming is the process of homogenizing a magnet
There are two types of shimming: [read slide]
On the right, an image I found the a paper that illustrates the MRI image of a hand without shimming, with passive only, and with both passive and active shimming.

## Slide 26: MFC 8246: Small Bore Mapper
Small bore mapper for fast and accurate magnet mapping without manual operations
It integrates a motorized stage and probe holder:
  360-degree rotation
  74.5mm vertical displacement
Measurement of the magnetic field with NMR
Configuration of either single probe or probe array
3. Magnetic Field Camera
Notes: It is in this context that our camera fits in middle: the flange makes the interface with the bore (adaptable fixation points) Small fixation bar are present between flange and motion stage to adjust the distance to magnet

## Slide 27: MFC 8246: Probes and Support
Hall probe
NMR Probe Head (PWP1526)
PWP 1526
Sample
5 mm
Hall probe
NMR Probe Head
with sample 5 mm
off-axis mapping
PWP 1426
MFC9146 Probe-Array
MFC9146
Fields are measured from 38mT up to 30T
The accuracy is <5ppm
High resolution down to 10ppb
3. Magnetic Field Camera
Notes: 1526 -> integrated component in chip
1426 -> discrete components
The timing of the mapping will depend on the probe choice (single or array) and number of measurement points:
e.g.: single probe15 planes (Z axel, 0 to 74.5) and 12 points (angular 0° to 360 °) – 30 minutes / probe array -> less than 5 minutes

## Slide 28: MFC 8246: Angular Motor
Motor for angular displacement:
  Non magnetic
  Has a travel range of 360°
  Resolution of 10μ°
  The motor can operate effectively in any orientation, whether it's positioned vertically, horizontally, or in any gravitational direction
Two motors for vertical displacement
  Non-magnetic
  74.5mm travel length
  4.5 μm step length
3. Magnetic Field Camera
Notes: For the angular motion we have a single…
Vertically in a spectrometer, or horizontally in an MRI, or any angle

## Slide 29: MFC 8246: Summarize
3. Magnetic Field Camera
Pulsed wave NMR magnetometer
Design to map small bore such as spectrometers, or small-bore MRI
Motorization
  accurate probe positioning
  performs the mapping without user operation
Successfully assessed at 6T, 9T, 11T and 14T.
Notes: To summarize, this magnetic field camera is a…

## Slide 30: Outlook
NMR simulation software:
  Depending on the customers' magnetic field, we simulate the NMR signal that our probes will obtain
Notes: As an outlook for this presentation, we are currently working on a software that could, based on simulated magnetic field from our customers, simulate the NMR signal that our probes will obtain.
This will allow us to provide the best solutions for our customers

## Slide 31
Thank You
Notes: This concludes

## Slide 32
Backup

## Slide 33: Shimming effect
Effect of active shimming on single voxel spectroscopy in the brain
  Inhomogeneities broaden the resonance peak of the metabolites
3. Magnetic Field Camera