## Slide 1: Cryogenic Tests of Electronic Components and Sensors for Superconducting Magnet Instrumentation
Vincenzo Di Capua, Marco Buzio, Lucio Fiscarelli, Unai Martinez
10/10/2024

## Slide 2: Table of contents
  Quench localization
  Quench localization methods
  Electronics for quench localization
  Needs
  Current limitations
  Embedded detection and controlling electronics
  Cryogenic electronics
  Selected components
  Evaluation board
  Test plan and setup
  Results
  Future perspectives
  Conclusions
Motivation
Proposal

## Slide 3: Quench localization
  Identifying the starting location of a quench during superconducting magnets training is critical:
    To identify defects.
    To drive the design and the manufacturing.
    For protection systems.

## Slide 4: Quench localization methods
Electrical
    Passive: voltage taps to monitor the resistance.
    Active: AC signals to measure impedance, reflectometry, or stray capacitances.
Magnetic
    Array of induction coils.
    Array of Hall sensors.
Mechanical
    Acoustic emission sensors.
Thermal
    Fiber-optic sensors.
Voltage taps and induction coils are widely used for LTS magnets.
All others are considered promising for HTS magnets.

## Slide 5: Electronics for quench localization
  Needs
  Acquire a large number (>40) of signals.
  Signal conditioning (mV range).
  High sample rate (>50 kHz).
Current limitations
The number of signals.
Signal integrity.
Environmental noise.

## Slide 6: Needs
  Acquire a large number (>40) of signals.
    Internal: V taps, induction coils, Hall probes.
    External: Quench antenna, Hall probes.
  Signal conditioning (mV range).
    Multi-stage amplifiers to increase the SNR.
    Antialiasing filtering.
  High sample rate (>50 kHz).
    To capture the full dynamic of the acquired signals.
Custom design of amplifier cards and crates for Quench Antenna.
Commercial acquisition system with real-time processing.

## Slide 7: Current limitations
The number of signals.
  Internal: Number of signals that are possible to extract from the cryostat.
  External: Acquisition channels.
Signal integrity.
  Due to several interconnections.
Environmental noise.
  Captured by sensors and cables.
Quench Antenna conditioning and acquisition rack.
sensor
Amplifier chassis
ADC1
ADC2
Cryostat (300/4.5/1.9 K)
Measurement hall (300 K)

## Slide 8: Embedded detection and conditioning electronics
  One Approach to mitigate most of these limitations would be to embed the whole chain into the sensors and/or into the cryostat:
    Signal conditioning:
      Shorter distances.
      Fewer interconnections.
    Digitalization:
      Reduced impact of external electromagnetic disturbances.
      We need to extract only a few digital signals.
sensor
ADC
conditioning
sensor
Amplifiers chassis
ADC1
ADC2
Cryostat (300/4.5/1.9 K)
Measurement hall (300 K)
Cryostat (300/4.5/1.9 K)
Digital readout
(300 K)

## Slide 9: Cryogenic electronics
  A crucial aspect to be investigated is to embed sensors and electronics in the superconducting magnets’ cryostats.
  A vast literature on using commercial microphones, Hall probes, and conditioning electronics at cryogenic temperatures (4.5 and 1.9 Keven if only in dry conditions)[1,2,3,4]
  Each component family and each manufacturer must be evaluated individually.
Hall probe array for cryogenic environments.
[1] Marchevsky, M. (2022). Understanding training in superconducting accelerator magnets using acoustic emission diagnostics. arXiv preprint arXiv:2203.08871.
[2] D. Conway Lamb, et All; An FPGA-based instrumentation platform for use at deep cryogenic temperatures. Rev. Sci. Instrum. 1 January 2016; 87 (1): 014701.
[3] Song, M., et All,(2018, March). Evaluation of commercial-off-the-shelf (COTS) electronics for extreme cold environments. In 2018 IEEE Aerospace Conference (pp. 1-12). IEEE.
[4] Valiente-Blanco, et All (2013). Characterization of commercial-off-the-shelf electronic components at cryogenic temperatures. Instruments and Experimental Techniques, 56, 665-671.
Hall probe sensors already in use on racetrack magnets (left) (credits: Carlo Petrone). Piezoelectric cryogenic sensor with embedded amplifier (right).

## Slide 10: Selected components
  The components were selected with the final goal of building a complete acquisition chain.
    Components have been chosen by taking inspiration from the existing literature.
    The test of digital components (FPGA, MCUs) has been postponed to the next test campaign.

## Slide 11: Evaluation board
  Custom PCB has been designed to host all the devices under test and provide a suitable interface for the signal cable and the 40 available feed-throughs.
  Test points have been added to simplify the debugging in case of faults.
  Form factor tailored to the test head.
  Standard industrial processes, FR4 dielectric, and ENIG surface finish.
PCB for 1st campaign of tests of sensor and electronic components at cryogenic temperature.

## Slide 12: Test plan and setup
  All the selected components have been tested at :
    300 K (room temperature)
    77 K (liquid nitrogen)
    4.5 K (liquid helium)
    1.9 K (superfluid liquid helium)
  Each test has been repeated twice:
    Immediately after power on.
    Check for performance stability.
  After each test, the PCB has always been brought back to room temperature.
Assembled test head (left), liquid nitrogen tests (right).
setup for the test in liquid Helium at 1.9 K.

## Slide 13: Operational Amplifier
  configured in non-inverting mode.
  DC gain of 10.
  Test (SNR evaluation):
    1.1 V AC input at 100 Hz.
    output acquired at 50 kHz.
    Four parameters Sinfit.

## Slide 14: MEMS microphones
  Acoustic vibrations are provided by a piezoelectric transducer.
  Test the full bandwidth of the sensor.

## Slide 15: Multiplexer
  In-out resistance was evaluated for two of the inputs of the multiplexer.
    Liquid nitrogen performance comparable with room temperature.
    Increase in liquid helium by a factor of almost 2, with a negligible difference between 4.5 K and 1.9 K.
  Component suitable for multiplexing signals at cryogenic temperatures.

## Slide 16: Future perspectives
Commercial MEMS microphones (left) and PCB board for integration in the magnet structure (right).
PCB for 2nd campaign of tests of sensor and electronic components at cryogenic temperature.
Test insert for critical current test on REBCO tapes.
  A second test campaign is foreseen.
    ADC, microcontrollers, and FPGA.
  A small PCB for integration of MEMS sensors on magnet structure has been developed and characterized.
  Acoustic emission during critical current tests of REBCO tapes is currently under study.
    Case study with only one MEMS sensor up to 15 T.

## Slide 17: Conclusions
Quench detection and localization are crucial for optimizing the design and manufacturing of superconducting magnets.
Typically, a large number of small signals (greater than 40, each a few millivolts) need to be acquired and sometimes extracted from the cryostat.
Using cryogenic electronics for signal conditioning and acquisition can improve the signal-to-noise ratio (SNR).
Individual tests are necessary to verify the functionality and characterize electronic components beyond the manufacturer’s specifications.
We successfully tested the signal conditioning chain's essential components and established a stable setup and procedure.
Future tests are planned to develop a complete cryogenic acquisition chain.
The tested MEMS sensors have already been deployed with success for cryogenic applications.

## Slide 18
Thanks for the attention!