Alvra - Short Pulse (Minimum) Requirements
Alvra - Short Pulse (Minimum) Requirements
Requirements in terms on pulse energy, stability, and transverse properties depending on which experimental method you employ.
Same requirements for all photon energies (likely in the range of 7-12 keV).
	Pink Beam Experiments	Monochromatic Beam Experiments
Pulse energy	> 400 uJ (minimum)	> 600 uJ (minimum)
Stability	Pointing / Jitter	Pointing / Jitter + Central energy
Bandwidth	Not Crucial	< 0.3% ideal
Pulse duration	<20 fs FWHM	<20 fs FWHM

---
Alvra - Beam Quality Measurements: Position jitter
Alvra - Beam Quality Measurements: Position jitter

---
Alvra - Short Pulse Measurements
Alvra - Short Pulse Measurements

---
Alvra - Effect of the dechirper
Alvra - Effect of the dechirper
Dechirper out (40 fs fwhm)	SiN (2 um)	68.8 fs
	water (~ 7 um)	60.6 fs
Dechirper in (20 fs fwhm)	SiN (2 um)	60.1 fs
	water (~ 7 um)	51.4 fs
Rise time on SiN always longer than on water (different excitation mechanisms)
Observed a change of ~ 10% in pulse duration when using the short pulse mode (dechirper in)

---
Alvra - Summary: General Experience with Short Pulses
Alvra - Summary: General Experience with Short Pulses
- We have been able to demonstrate 30-35 fs FWHM in Alvra → X-ray could not have been longer then 20 fs FHWM from our calculations (reported 38 fs FWHM e-bunch)
- We unfortunately see very little correlation between the measured FWHM in our experiment and the reported pulse duration (especially from electron beam measurements) → Inconsistent measurements
- We also don’t have a strong correlation between the dechirper and the measured IRF. After using it 4+ times, we see a (small) effect for the first time on our last beamtime
- The short pulses we have gotten with the “optimized” compression have been more reliable. They are always shorter and have more flux. We would like to explore this more and be able to change compression on demand.
- Joint commissioning+MD time to define the parameter space would be highly appreciated (horns lasing, compression, pulse duration measurements, pulse energy)