## Slide 1: SwissFEL Seed Lasers Status
A. Trisorio: on behalf of the Gun Laser Group: Paul Scherrer Institute
12.03.2024

## Slide 2: Seed laser facility overview
Castor / Seed laser 1
Laser Transfer Line 1
Laser Transfer Line 2
1st magnetic chicane
2nd magnetic chicane
2nd Modulator (U200)
1st Modulator (U200)
10 m
Laser Setup 1 – IR / UV
Laser Setup 2 - UV
Pollux / Seed laser 2
Installed & Commissioned

## Slide 3: Seed Lasers - architecture
Common oscillator: Gecco-Ultra (Novanta)
RF synchronization
Synchronization quality improved: <16 fs RMS over 10Hz – 10MHz
Further R&D ongoing to improve timing jitter/drift
Twins amplifiers: Pulsar (Amplitude)
High pulse energy stability < 0.5 % rms (multiple pump lasers)
Robust pump laser units (Centurion – Quantel)
Wavelength tunning capability (Dazzler + Mazzler)
Active stabilization of the compressed pulse arrival time (in the lab) via
Balanced Optical Correlators (BOC) + feedback loop
Compressors moved to the tunnel optical tables after the
Optical transfer lines (avoid nonlinear effects)

## Slide 4: Seed Lasers – few photos
Laser 1 – Castor
(installed December 2021)
Laser 2 – Pollux
(installed December 2022)

## Slide 5: Seed Lasers – tunnel injections setups
2nd Modulator
Z=410
2nd Laser TL
1st Modulator
Z=350 m
1st Laser TL

## Slide 6: Seed Lasers – injection setup 1
Seeding of modulator 1 possible with:
T=60 fs FWHM – E≤5 mJ @ λ=793 nm
OR with T=160 fs FWHM – E≤0.45 mJ @ 261 nm<λ<272 nm

## Slide 7: Seed Lasers – injection setup 2
We have upgraded the optical setup to be able to seed:
T=160 fs FWHM – E≤0.35 mJ @ 261 nm<λ<272 nm

## Slide 8: Changes compared to previous run IR
Check and adapt collimation of IR beam into the compressor
- Exchange small compressor grating: golden layer showed dark/bright structures
Compressor output characterized
IR Pulse temporal profile – T=61 fs FWHM (FL=51 fs)
IR Spectrum
λ=792.6 nm /Δλ=19.5 nm FWHM

## Slide 9: Changes compared to previous run UV
Removed IR telescope before UV generation
Cleaned all optical elements on the setup and check stress
Replaced 1 damaged UV optic
Damaged in-coupling viewport was exchanged
Laser/Electron beam profile Diagnostics (Juranic):
New objective to increase the resolution (8 um with YAG screens)

## Slide 10: Modulator 1 – 8500 mm lens
Virtual path: Measured with Wincam.
Modulator Center
down stream
up stream
Modulator Input
Modulator Output

## Slide 11: Modulator 1 – 8500 mm lens
Virtual path: Measured with Wincam.
Modulator Center
down stream
up stream
Modulator Input
Modulator Output
Good UV beam quality through the Modulator

## Slide 12: Setup for M2 measurement
Focal length 1000 mm
Knife Edge on Newport MFA-PPA stage
Z-Translation on Owis 300 mm linear stage
Energy measurement with high resolution energy probe
(Wieser Photodiodes)
Input beam
Moving knife
edge
Pulse energy
detection
Z - Translation

## Slide 13: Modulator 1 – before change – full beam
Scaled to 8500mm lens:
zr: 903 mm
Radius w0: 305 um @1/e2
Lens plane
16.10.2023

## Slide 14: Modulator 1 – before change – full beam
Scaled to 8500mm lens:
zr: 397 mm
Radius w0: 194 um @1/e2
SSL-LCAM-C1103
imaging plane around WP
size not calibrated
Lens plane
SSL-LCAM-C1104
Modulator plane

## Slide 15: Modulator 1 – 10 mm iris
Scaled to 8500mm lens:
zr: 549 mm
Radius w0: 228 um @1/e2
Lens plane
SSL-LCAM-C1103
imaging plane around WP
size not calibrated
SSL-LCAM-C1104
Modulator plane

## Slide 16: Modulator 1 – 8 mm iris
Scaled to 8500mm lens:
zr: 759 mm
Radius w0: 269 um @1/e2
SSL-LCAM-C1103
imaging plane around WP
size not calibrated
Lens plane
SSL-LCAM-C1104
Modulator plane

## Slide 17: Modulator 1 – UV Pulse duration
UV pulse duration of 163 fs FWHM
(Assumed 132 fs IR pulse due to dispersion)
EIR=5.4 mJ input /EUV= 0.7 mJ
UV Spectrum
λ=264.7 nm /Δλ=1 nm FWHM
We are well capable of delivering expected UV beam parameters in modulator 1

## Slide 18: Modulator 2 – clipped beam as used in November 2023
Scaled to 8500mm lens:
zr: 910 mm
w0: 301 um @1/e2

## Slide 19: Modulator 2 – full beam as used in November 2023
Scaled to 8500mm lens:
zr: 680 mm
w0: 264 um @1/e2
More time required to improve and fully characterize seed laser 2 (@next shutdown?)

## Slide 20: Conclusions
We have a commissioned Seed laser facility ready for MD shifts
We identified the source of the UV beam spatial structures: mainly IR compressor grating
Seed laser 1 UV beam fully characterized
The laser beam properties / parameters are well inside the specified ones
For Seed laser 2, we may improve the parameters with some time to work on the setup
What will be required for the Seed Lasers (and in what time frame)
to have an operative facility for experiments?

## Slide 21: Wir schaffen Wissen – heute für morgen
Our thanks go to:
Timing & Synchronization, Diagnostics, Beam Dynamics & Operation groups
Control’s supporters
S. Neppl, A. Cavalieri

## Slide 22: Seed Lasers - few numbers
Seed Lasers performances in a nutshell
Laser wavelength (nm) | 800 / 266
Pulse duration on target FWHM (fs) IR/UV | 80 – 600 / 200
Pulse energy on target (mJ) IR/UV | 5 /0.2
Pulse energy stability on target ( % rms over 1000 shots) | <0.35 / <1.3
Beam diameter on target (um) | X=433 / Y=394 /
Relative pointing stability on modulator 1 (% of beam zize) | <1.1% over 30 min
Oscillator integrated jitter (fs rms) Over 10 Hz- 10MHz bandwidth | 16
Amplifier arrival time jitter with respect to oscillator (fs rms) over 1 min | <2.7
Relative time drift UV seed laser with respect to electrons (fs over 7h) | <36
November 2023, laser 800 nm
Seed Laser Synch. quality improved a lot !
Joint effort with T&S group

## Slide 23: Seed Lasers Focusing geometry
Stage 1 - UV
Stage 2 - UV
THG
output
Mod 1
center
Mod 2
center
VA
HWP + TFP
F=8.5 m
UVFS
HWP
Quartz viewport
3 mm/uncoated
Motorized iris
THG
output
F=8.5 m
UVFS
HWP
Quartz viewport
3 mm/uncoated
2 lenses telescope
R=2/3

## Slide 24: Seed Lasers – UV generation
Deep-UV generation stage mounted and commissioned: required for EEHG
>400 uJ UV available in the tunning range, <1.3% RMS stability over 500 shots
Chart: UV Spectra

## Slide 25: Seed Laser 1 Upgrade December 2022
We have the compression in tunnel ->up to 5mJ/100fs FWHM @ 800 nm
We installed the UV generation stage