Machine Report Week 21/2024
Thomas Schietinger
SwissFEL Exchange Meeting, 27 May 2024

---
Introduction
Calendar

---
Photon (pulse) energies
Aramis
(unfortunately, no data available
from archiver directly)
9.5–10 keV for
Bernina, ~900 µJ
 ● Photon energy (keV)
 ● Photon pulse energy (µJ)
Mon,20-May, 08:00 – Mon, 27-May, 08:00
9.5–10 keV for
Bernina, ~900 µJ
Setup
7.1 keV
for Alvra
Interruptions
from SINSB02
Athos
RF intervention
for SINSB02
500 and ~400 eV for Maloja,
2.5–3.0 mJ
Machine
development
Setup

---
Setup
Aramis
- Standard setup for 9.5 keV. Focus on
  - optimization of bandwidth
  - mitigation of beam losses
- Preparation of 7.1 keV (Alvra)
Athos
- Preparation of 400 eV and 500 eV, both with linear horizontal polarization
- Maloja takes beam parasitically during setup

---
Overall machine status
The Aramis branch suffers from two issues right now (possibly related):
Undulator alignment:
- The first four undulators are not contributing to lasing due to misalignment.
- A beam-based alignment (BBA) was attempted by Sven Reiche. It helps but cannot solve all problems.
- Next iteration planned for June 11/12 (align undulators on BBA orbit, FEL optimization on this orbit)
- Some machine components, in particular at the beginning of the undulator line, require a physical realignment in the tunnel to fulfill pointing requirements. This needs careful planning and many working hours from the survey group (which is superbusy with SLS-2!)
- Overall, we can achieve about 75% of maximum performance in Aramis
  (i.e., 0.9 mJ at 9.5 keV, where we would expect 1.2 mJ)
Beam losses:
- Increased losses since about three weeks
  (i.e., already before the BBA)
- Root cause not found yet. Halo?
- Last week, the losses appeared to increase exponentially
  (losses get worse with increased integrated beam current?)
exponential increase of
undulator radiation dose?

---
Issues last week
Causing machine downtime:
Interruptions from RF station SINSB02 (Monday and Tuesday): frequent interlocks from transformer temperature and JUMO communication error. Intervention by RF team on Tuesday morning solves both issues (exchange of HVPS, debugging of JUMO). (3 hours downtime)
Other issues (no machine downtime)
- Losses in Aramis undulator. A few DRM (dose rate monitor) resets were needed.
  - Many knobs tried (matching quads, collimators, opening of first undulators, gun laser parameters, gun solenoid, etc.) – with limited effect
  - Best knob was orbit in energy collimator (shift orbit target of energy feedback away from zero)
- S20SY02-MKDC050 has too high resistance (gives alarm) (Tuesday morning). Bridged, was fixed on Monday morning
- WCM failures (Tuesday, 15:00, Sunday, 6:30), MPS stops the machine. Requires restart of IOC.
- Failure of first solenoid magnet power supply (the infamous bucking coil) (Wednesday, 11:00)
- Mizar timing jump → sudden broadening of Athos spectrum (Thursday, about 18:00)
- Instable beam (Aramis) due to feedback stopping due to a screen obstacle (from a screen removed a long time ago…) (Friday, 8:00)