## Slide 1: FRESCA Test Facility CERN Sc cable test stations
J. Fleiter, F. Girardot, A. Vande Craen and T. Boutboul
Contributions from many CERN colleagues in different departments
5th Superconducting Magnet Test Facility Workshop March 29 – April 4, 2025, Bad Zurzach, Switzerland

## Slide 2: Outline
FRESCA1
FRESCA2
Fabrication and installation of FRESCA2
Electrical circuits
Cryogenic distribution
Overview of typical measurements in FRESCA

## Slide 3: Introduction to FRESCA
FRESCA : Facility for Reception of Superconducting Cables (High Energy collider application)
Electrical characterization of Sc cables in LHe and in transverse field of up to 13 T, located @ CERN in b 163
  FRESCA 1:
    in operation @ CERN since 1998
    T=1.9-4.3 K
    Field of up to 9.6 T
    Used for qualification of LHC and HL LHC cables
  FRESCA 2:
    Installed @ CERN, being commissioned
    T=1.9-4.3 K in LHe, possible higher Temp in GHe
    operating field of up to 13 T
    Will be use for qualification of HFM and R&D cables

## Slide 4: Type of measurements
Rutherford cables (Nb-Ti, Nb3Sn), round cables (Nb-Ti, MgB2,HTS), HTS stacks , Roebel cable…
UI-measurements (B, B-direction, T and dI/dt) ⇒Ic, Iq, n-value
UB-measurements (I, B-direction, T and dB/dt) ⇒flux jump stability
RRR measurement (as function of B if needed)
Measurement of the self field along the cable ⇒transport current distribution among the strands, BICCs
Measurement of splice/joint resistances (accuracy better than 0.05 nΩ)
Measurement of (Minimum) Quench Energy (as function of B, I, T, heater position, heat pulse duration)
Measurement of quench velocity (as function of B, T)
More details in next slides
A.P. Verweij, Workshop on test facilities and measurement equipment needed for the LHC exploitation , 12 April 2006

## Slide 5: CERN strand and cable Test Facility

## Slide 6
CERN strand and cable Test Facility
  Test stations 1-4: Strand Ic (12-15 T, 1.9-4.3 K)
  Test station 5: cable magnetization (1.8 T. 1.9-4.3 K)
  Test station 6: VSM (10 T, Variable Temp.)
  Test station 7: RRR and inter-strand resistance
  FRESCA1: cable Ic test station (9.6 T 1.9-4.3 K)
  FRESCA2: cable Ic test station (13 T, 1.9-4.3 K)
  Upgrade of Test station 1: Strand Ic (20 T, 1.9-4.3 K, then VTI)
  Upgrade of test station 5 : Strand Ic (split pair 10 T, 1.9-4.3 K, then VTI)
  Upgrade
VSM
RRR
Magnetization
FRESCA1
FRESCA2
Ic strands 1-4
Cold Box
5 g/s
Buffer
6 kLHe

## Slide 7: FRESCA1
Built in 1995-1998, operational since then
Magnet:
  88 mm dipole up to 10.1T (at 1.9 K, ~12 kA),
  Operating field of 9.6 T
  Magnet housed in an independently cooled cryostat.
  Current leads of 18 kA
  Power supply of 16 kA 6V 1Q
  Dump resistor of 40 mOhm
Sample
  Housed in a second cryostat (called Inner Cryostat ID ~70 mm)
  Test volume (homo. field) 600 mm x 70 mm ∅
  Free aperture in Sample holder for Nb-Ti and Nb3Sn cables 35x35 mm2
  Sample current with external 32 kA PC or internal Sc transformer (~50 kA).
A.P. Verweij, 1.9 K Test Facility for the Reception of the Superconducting Cables for the LHC, IEEE Trans. Appl. Supercond. , VOL. 9, NO. 2, JUNE 1999
D. Leroy, Design and Manufacture of a Large-Bore 10 T Superconducting Dipole for the CERN Cable Test Facility, IEEE Trans. Appl. Supercond. 10, 1 (2000) pp.178-82

## Slide 8: FRESCA1 Test Facility in 1999

## Slide 9: FRESCA2
Being commissioned in 2025
Magnet:
  100 mm aperture dipole up to 14.6 T (at 1.9 K, ~12 kA)
  Operating field of 13 T, 10.6 kA
  Magnet housed in an independently cooled cryostat.
  13 kA HTS Current leads of LHC
  Power supply of 16 kA 6V
  Dump resistor of 80 mOhm
Sample
  Housed in a second cryostat (called Inner Cryostat ID of ~82 mm)
  Test volume (homo. field) 650 mm x 82 mm ∅
  Sample current with external 32 kA PC or internal s.c. transformer (>50 kA).

## Slide 10: FRESCA2 Inner / Outer Cryostats

## Slide 11
Outer cryostat assembly
Inner cryostat assembly
Magnet assembly
Inner
He vessel
Lambda plate
Magnet current leads
Inner
thermal screen
Outer
thermal screen
Outer
He vessel
Inner
vacuum vessel
Outer
vacuum vessel
Alignment system
Radiation screen assembly
FRESCA2 cryostat
FRESCA2 Assembly
Slide from B. Riffaud, on behalf of CERN EN-MME Design Team

## Slide 12: FRESCA2 Magnet Test @SM18
Bloc coil with flared ends
100 mm Ø free aperture
Magnet tested in SM18
  14.6T @1.9 K
  13.9T @ 4.3 K
G. Willering et al., "Tests of the FRESCA2 100 mm Bore Nb3Sn Block-Coil Magnet to a Record Field of 14.6 T," IEEE Trans. Appl. Supercond. , vol. 29, no. 5, pp. 1-6, Aug. 2019
CEA–CERN Block-Type Dipole Magnet for Cable Testing: FRESCA2. In: Schoerling, D., Zlobin, A. (eds) Nb3Sn Accelerator Magnets. Particle Acceleration and Detection. Springer, Cham.

## Slide 13: Transfer of torque magnet/sample
13
Top part
Rigid fixation to the lambda plate
Sample holder
Torx angular positioning with the inner helium vessel
Inner helium vessel
Inner vacuum vessel
Centering (rotation free) with the micrometric xy table
Bottom part
Micrometric angular indexer
Torx angular positioning
Inner helium vessel cane
Sample holder
Rigid fixation to the magnet
Magnetic torque between sample under test and magnet needs to be transferred via the holder and inner cryostat
Top and bottom of holder locked in rotation

## Slide 14: Integration of FRESCA2
Installed in the existing pit, no civil engineering work
Half of the top plate shall be free of protruding components
Cryostat position:
  Top flange of inner cryostats (F1 and F2) shall be at the same heights
  The centre of FRESCA 2 cryostat at same distance from wall as FRESCA 1

## Slide 15: FRESCA2 pit constraints
Small clearance cryostat /pit’s wall
Limited accessibility during installation
  No access to cryostat legs once cryostat is inside the pit
  Tight alignment of components required
  Welding of cryogenic connection were quite challenging

## Slide 16: FRESCA 2 Cryostat fabrication
Construction of FRESCA2 Cryostat by CERN main workshop, E. Rigutto (EN) on behalf of the project (EN)
Performed @ CERN

## Slide 17: Assembly of FRESCA 2
Assembly of FRESCA2 Cryostat by CERN main workshop (EN) in collaboration with Sc team

## Slide 18: Installation of FRESCA 2 in b163

## Slide 19: FRESCA2 Test Facility in 2024

## Slide 20: FRESCA Field profile
Dipole magnets
  FRESCA1 warm bore of 70 mm , 9.6 T
  FRESCA2 warm bore of 82 mm, 13 T
  3% homogeneity over ~600/650 mm
Cable sample splices placed out of field

## Slide 21: FRESCA Sample Insert
FRESCA1 and FRESCA2 shares the same inserts:
  2x 32 kA vapor cooled leads
  1x Sc transformer 50 kA
Sample insert could be rotated by 90º when sample is at 4.3 K or 1.9 K
  Applied field Perp or // to cable face
  Change of field direction by polarity invertor
Hydraulic jack lifting the insert to unlock the antirotating system
Leak tightness maintained during rotation

## Slide 22: Sample Holder of FRESCA1
  Standard Rutherford cable holder:
    Welded stainless steel tube structure
    OD 69 mm
    Inner cavity 35x35 mm
    Pre-stress via M16 screw
  High pressure holder
    Based on bladder and Keys
    Al shell
  Cable mounted on plate (low Lorentz force)
    Cable wrapped with Kapton
    Cable assembly inserted in G10 structure
B. Bordini, Critical Current Measurements of High-Jc Nb3Sn Rutherford Cables under Transverse CompressionIEEE Trans. Appl. Supercond. vol. 24, no. 3, 2014
P. Ebermann et al Irreversible degradation of Nb3Sn Rutherford cables due to transverse compressive stress at room temperature 2018 Supercond. Sci. Technol. 31 065009

## Slide 23: Electrical Distribution
Power converters
  Three 16 kA -6 V power converters each equipped with DCCT and FGC2
    16 kA for the magnet
    32 kA for the sample
  Sample could also be operated from a 50 kA Sc Transformer (PC 200 A, 4Q )
FRESCA magnet polarity inverter
  allow to change the direction of the field applied to the cable samples.
  Interlocks switches to ensure it is in one position or in the other.
FRESCA magnet selection switch
  allow to energize from a single 16 kA PC either FRESCA1 or FRESCA2 magnet
  Interlocks switches to ensure that the switch is in one position or in the other.
FRESCA sample quick disconnect
  Use to connect quickly the 32 kA PC to the sample inserted either in FRESCA1 or in FRESCA2
  Interlocks switches provided to ensure that the switch is in one position or in the other.

## Slide 24: Station 32 kA Quick connect
Allow with one PC 32 kA to operate either:
  Samples FRESCA1
  On short circuit
  Samples in FRESCA2
Water cooled cables attached to a slidable clamp (pneumatic clamping)
Time of swapping ~3 mins
Quick connect 1999-2019
Quick connect 2020=>

## Slide 25: Magnet Protection
Redundant Quench Detection System (developed at CERN)
Fast Mechanical switch, opening in less than 40 ms
Dump resistor 40 mOhm (F1) and 80 mOhm (F2)
No quench heaters

## Slide 26: Data acquisition system
DAQ shared between FRESCA1 and FRESCA2
DAQ consists of commercial/CERN developed test equipment coupled to a PXI workstation.
  High accuracy low Voltage (nV) measurements (~10 Hz / channel )
    8 x Keithley 2182A, inputs channels selected via matrix
  High frequency Voltage measurements
  (100 kHz)
    Gen3i transient recorder
    uQDS CERN developed for quench detection (see J. Steckert presentation, Thursday)

## Slide 27: Cryogenic operation of FRESCA2
Operation performed by an automate that controls:
  Purging Cool-down (300 K to 4.3 K and 4 K 1.9 K)
  Nominal operation at 4.3 K or 1.9 K
  Recovery of nominal operation after quench
  Warm up (1.9 K to 4.3 K)
  Drain and warm up (4.3 K to 300 K)
ΔT<50 K for Magnet during cool down and warm up
Maximum cool down/warm up time
Cryo. OK signals to QPS
        LHe level at its nominal
        Temp of magnet and sample leads
        Magnet bath temp. and pressure

## Slide 28: Cryogenic distribution system
One LHe cold box ~5g/s
One buffer 6 kLHe
Few valve boxes, many GHe racks
9 test stations (2 cables, 7 strands)
Recovery GHe system

## Slide 29: Overview of typical measurements
Cable Ic measurements LTS/HTS
Cable Splice resistance
Normal zone propagation velocity
MQE and current distribution

## Slide 30: Cable Ic measurements NbTi /Nb3Sn
Qualification of LHC cables cable Ic measured in FRESCA (1.9 K/ 4.3 K) and at BNL (4.3K)
Qualification of HL-LHC cables and FCM cables
Ic degradation of Nb3Sn cables
Transverse stress applied to cables:
  At room temperature
  In situ with a specific holder
Critical Current Measurements of the Main LHC Superconducting Cables IEEE Trans. Appl. Supercond. VOL. 17, NO. 5, 2007
P. Ebermann et al Irreversible degradation of Nb3Sn Rutherford cables due to transverse compressive stress at room temperature, Supercond. Sci. Technol. 31 065009, 2018
L. Bottura, “Strand and cable R&D for fast cycled magnets at CERN,” IEEE Trans. Appl. Supercond., vol. 21, no. 3, pp. 2354–2358, Jun. 2011
J.E. Jean-Eudes Duvauchelle, Critical Current Measurements Under Transverse Pressure of a Nb3Sn Rutherford Cable Based on 1 mm RRP Wires,Appl. Supercond., vol. 28, 2018
J. Fleiter Quench Propagation in Nb3Sn Rutherford Cables for the Hi-Lumi Quadrupole Magnets", IEEE Trans. Appl. Supercond, vol.25, no.3, pp.1-4, 2015.

## Slide 31: Cable Ic measurements MgB2 / HTS
MgB2 cables of Sc Link 4 K
  Up to 22 kA in s.f.
REBCO ROEBEL 4K
15 kA sf, 12 kA 8 T
CORC cable 4K
Twisted Pair REBCO 4K
K. Konstantopoulou, Electro-Magnetic Performance of MgB2 Cables for the High Current Transmission Lines at CERN, Appl. Supercond. ., VOL. 29, NO. 5, 2019
T. Mulder Performance Test of an 8 kA @ 10-T 4.2-K ReBCO-CORC Cable,, IEEE Trans. Appl. Supercond. VOL. 26, NO. 4, JUNE 2016
A. Ballarino First tests of twisted-pair HTS 1 kA range cables for use in superconducting links Physics Procedia Vol. 36, 2012
J. Fleiter Electrical characterization of REBCO Roebel cables. Supercond. Sci. Technol. 26 065014 2013

## Slide 32: Cable Splice resistance
LHC Nb-Ti cables
MgB2 cables Sc Link
The Electrical Resistance of Rutherford-Type Superconducting Cable Splices IEEE Trans. Appl. Supercond. , VOL. 25, NO. 3, 2015
S. Giannelli Investigation of Splice Resistances of High-Current MgB2 Cables Operated in Liquid and Helium Gas IEEE Trans. Appl. Supercond. , VOL. 27, NO. 4, AUGUST 2017

## Slide 33: Normal zone propagation velocity
J. Fleiter Quench Propagation in Nb3Sn Rutherford Cables for the Hi-Lumi Quadrupole Magnets", IEEE Trans. Appl. Supercond, vol.25, no.3, pp.1-4, 2015.
NZPV measured by time of flight
Quench triggered by Heater
Transverse delay quench

## Slide 34: MQE and current distribution
MQE on Nb-Ti Nb3Sn cables
Current distribution
A. Verweij. , Experimental results of current distribution in Rutherford-type LHC cables Cryogenics 40 (2000) 663-70
W. M. de Rapper et al., Thermal Stability at 1.9 K and 4.3 K of Nb3Sn Cables for Quadrupole Magnets for the LHC Upgrade IEEE Trans. Appl. Supercond. , vol. 23, no. 3, 2013
G. Willering et al., Stability of Nb-Ti Rutherford Cables Exhibiting Different Contact Resistances IEEE Trans. Appl. Supercond. , vol. 18, no. 2, 2008

## Slide 35: Next steps
Electrical commissioning of FRESCA2 May-June 2025
First test of Sc cable Q3 2025
Develop a VTI operated with GHe

## Slide 36: Conclusions
FRESCA:
  A versatile cable test station
  Mandatory tool to qualify Sc cable for accelerators applications
  In operation since 25 years
  Installation of FRESCA2 completed
  Cryogenic Commissioning Completed
  Electrical commissioning to be completed by mid 2025

## Slide 37: Thank you for your attention