## Slide 1: ATAR Sensor Development
Jessie Yang
University of Washington

## Slide 2: Current ATAR Design
For PIONEER, ATAR needs precise timing, position, and energy measurements.
Timing / position precision is largely achieved with a highly granular ATAR + thin silicon sensor layers.
  Total size is 2 cm x 2 cm x 6 mm such that pions and muons stop in target
  Granularity is achieved with 48 layers of sensors, each with 100 strips of 2 cm x 0.02 cm Layers will provide position information in two dimensions (z and either x or y). Consecutive layers will alternate between providing x and y information.
  120 um thick layers allow for response time.
Energy precision introduces technological challenges
  Must be sensitive enough to detect MIPs and reveal positron track
  Must also measure pion/muon energy for various cuts (e.g. dE/dx and prompt vs. delay cut)
  This is a dynamic range of around 100
ATAR Sketch
Strips
Dead material

## Slide 3
Incoming particle
Reverse bias is applied to amplify electric field in all regions
Intrinsic layer (undoped)
  LGADs have an additional, highly p-doped layer close to the N-junction
  Increases electric field in this region
  Avalanche effect enhances the signal for a given sensor thickness
  Gain is the multiplication factor by which a signal is amplified
  A sensors gain depends on the bias voltage applied and the incident angle of the incoming particle
Low-Gain Avalanche Diodes
Based on silicon PIN diode
  P-N junction separated by an intrinsic layer (undoped silicon)
  Reverse bias maintains electrostatic equillibrium until incoming particle deposits energy

## Slide 4: Test Beam Goals
Studying detector response to high energy deposits (100 MiPs)
Pulse shape features
Timing
Energy
Gain saturation?
Gain saturation dependencies
Incident angle
Incident energy
Detector production
Bias voltage

## Slide 5: Using CENPA’s Tandem Accelerator for PIONEER
Beam direction
Beam line
LGAD
SiMon
24” RBS Cahmber
Gold foil target wheel
Beta source tests

## Slide 6: Over 1000 Runs with 100,000+ Events Each
Sensor | Thickness (um) | Beam energies (MeV) | Bias voltages (V)
FBK W1 | 50 | 1.8, 3 | 50, 150, 250
FBK PIN W1 | 50 | 1.8, 3 | 50, 100, 400
FBK W8 | 100 | 1.8, 3, 4.5 | 100, 200, 250, 350
FBK PIN W8 | 100 | 3, 4.5, 6.5 | 200, 300, 400
FBK W8 (thinned) | 100 | 6.5, 8 | 100, 200, 250, 300
FBK W11 | 150 | 3.8, 5.5 | 200, 400, 600
FBK PIN W14 | 150 | 3.8, 5.5, 6.5 | 300, 500
BNL | 200 | 1.8, 3, 4.5, 5.5 | 200, 300, 400
scan of 10 different angles and sometimes through the backside
+
New code was developed to automate data taking (e.g. bias voltage and angle).
Huge thanks to the CENPA beam team for long hours, adaptability, and a consistent beam!
Beam energies chosen to study signals from both stopped and punch-through particles for each sensor thickness

## Slide 7: Trends from last years test beam
These factors affect charge density within the sensor which shields a given region from the amplified electric field, suppressing the measured gain.
Large bias will magnify the E field and thus the signal Great for positrons
Gain increases with angle
Gain decreases after some critical angle

## Slide 8: CENPA Test Beam 2024 Results
Gain = MPV of energy measurement of LGAD / PIN

## Slide 9: CENPA Test Beam 2024 Results
Gain = MPV of energy measurement of LGAD / PIN

## Slide 10: Comparing to 2023 Results (50 um sensors)
FBK W1, 1.8 MeV
FBK W1, 1.8 MeV
FBK sensors have more gain for MIPs but also more gain suppression for non-MIPs
  Even at the lowest biases/thinnest sensors/lowest energies, there is still significant gain suppression
Due to this suppression, the FBK boards have less gain for non-MIPs than HPK
The trends with respect to bias voltage and angle are still present

## Slide 11: Comparing to 2023 Results (50 um sensors)
FBK sensors have a shallower but more highly doped gain layer
FBK
Depth
HPK
Doping concentration
Depth
Doping concentration
Simultations and future testing will help us better understand the individual affects of both depth and doping concentration on gain saturation

## Slide 12: Understanding pulse shapes
Example Average Pulses: FBK W1, 50 um, 3 MeV
50 V
150 V
Variety of sensor thicknesses
  Thicker sensors need larger voltage biases for clean pulses
  Simulations will also be used to study multi-peak pulses and charge drift mechanisms within the sensors

## Slide 13: Received FBK Trench-Insulated LGADs
Production from AIDA 2020
Several devices, all 50um thick
  From the same wafer W10 (unsure about proprieties)
  Longest strips are 6mm with 100um pitch and other shorter strips
  Several 2x1 testing structures
Testing in progress
Slide curtesy of Dr. Simone Mazza

## Slide 14: TI-LGAD laser TCT testing
  Laser TCT testing of the device with the opening on the front
    6mm long strip, 100um pitch
  Observed a 3% first neighbor cross-talk
  “Long range” crosstalk around 1%
  Backside open (grid) for testing
Slide curtesy of Dr. Simone Mazza

## Slide 15: X-ray
  Tested another FBK TI-LGAD at the SSRL test beam
    Focused beam (BL 7-2)
  Shorter strips, 200um pitch
  For MIP the energy resolution is 10-15% as expected
Slide curtesy of Dr. Simone Mazza

## Slide 16: Plans
Continuing CENPA 2024 test beam analysis:
  Incident particles entering from the backside
  Quanitfy energy resolution and timing precision
  LGAD simulations
Next CENPA testing:
  Test FBK devices with deeper gain layer (from HGTD)
  Test HPK sensors with deep gain layer but varying doping concentration
We should receive mid-2025 a production of FBK TI-LGAD that are thicker (~100um active thickness)
  Strip geometry, pitch/length still to be decided
  Great option for the demonstrator
  See if it’s feasible to influence the gain layer doping to reduce doping and make the gain layer deeper
In the meantime continue to exercise thinning procedure
  Necessary for the demonstrator assembly

## Slide 17: Backup Slides

## Slide 18: Trends in Gain Suppression
50 um: as expected, greater saturation at larger voltages
100 um: unexpected, but beta source tests are similarly odd