## Slide 1: High-Density OPM-MEG
Joe Gibson
WOPM 2025
Joseph Gibson1, Holly Schofield1,2, Ewan Kennett1, Niall Holmes1,2, Ryan M. Hill1,2, Matthew J. Brookes1,2
1Sir Peter Mansfield Imaging Centre, School of Physics and Astronomy, University of Nottingham, Nottingham, United Kingdom
2Cerca Magnetics Limited, 7-8 Castlebridge Office Village, Kirtley Drive, Nottingham, United Kingdom

## Slide 2
SQUID system
OPM system
>300 Channels
<200 Channels

## Slide 3: Overview
High density OPM-MEG improves spatial resolution and sensitivity
Accurate array calibration is important for high density OPM-MEG
Increasing channel count is necessary to maintain high density coverage and reduce spatial leakage

## Slide 4: Moving closer to the scalp
Boto et al., 2016 10.1371/journal.pone.0157655
Iivanainen et al., 2021 10.1016/j.neuroimage.2021.118747
Sensitivity and spatial resolution improve closer to the scalp
40
-40
-4
4
Effective number of spatial components
200
80

## Slide 5: Moving to higher density
Boto et al., 2016 10.1371/journal.pone.0157655
Hill et al., 2025 10.1162/imag_a_00535
Sensitivity and spatial resolution improve at higher density

## Slide 6
Where we want to be
Where we are
???
Crosstalk
Coregistration
Sensor footprint
Heating
Interference rejection
Cost
Electronics

## Slide 7: Crosstalk
Interference effect between close OPMs
Can alter gains and sensitive orientations
With higher density arrays, we need a way to characterise and account for crosstalk

## Slide 8: Calibration Methods
Source localisation requires accurate knowledge of array geometry
More crosstalk = reduced source localisation accuracy

## Slide 9: Calibration Methods
Matrix coils used to create known fields (96 coils around the room)
Use these fields to derive positions, orientations and gains of the sensors
Iivanainen et al., 2022 10.3390/s22083059

## Slide 10: Calibration Methods
Create fields to represent 3 uniform fields and 5 gradient fields at different frequencies
, , , , , , ,
Measurement of uniform fields give orientation and gain
Measurement of gradient fields give position
Schofield et al., 2025 (in review)
Iivanainen et al., 2022 10.3390/s22083059
Hill et al., 2025 10.1162/imag_a_00535

## Slide 11: Calibration Methods
Residuals from CAD positions
1.3 ± 0.9 mm
Schofield et al., 2025 (in review)
Iivanainen et al., 2022 10.3390/s22083059
Hill et al., 2025 10.1162/imag_a_00535
Calibration gives accurate geometry
Short waveform (down to 1s)
Quick to solve (~1s)

## Slide 12
Hill et al., 2024 10.1162/imag_a_00112

## Slide 13: High Density Methods
High density helmets allow for changing array geometry
  Bespoke helmets from MRI for accurate coregistration
Sensors can be packed over relevant areas for different paradigms

## Slide 14: Spatial Resolution – pilot
Index
Little
Finger abduction paradigm
Localised to motor cortex
Different for two digits

## Slide 15: Spatial Resolution – pilot
Index
Little
Turn off random channels in analysis
Non linear relationship between fieldmap similarity and measured field
Increasing sensor density

## Slide 16
Hill et al., 2024 10.1162/imag_a_00112
It works! But…
Only image specific regions of the brain
Limited by sensor footprint

## Slide 17
Hill et al., 2024 10.1162/imag_a_00112

## Slide 18: 384 channel Methods
128 triaxial QuSpin OPMs (384 channels)
2 x Neuro-1 electronics
Synchronised by external clock signal (provided by signal generator)
Schofield et al., 2025 (in review)
Schofield et al., 2024 10.1162/imag_a_00283

## Slide 19: 384 channel Methods
Schofield et al., 2025 (in review)
Visual grating task
Gamma band Sensitivity
Phantom experiment
Localisation accuracy and crosstalk effects
Movie watching task
Spatial resolution

## Slide 20: 384 channel Results – Phantom
Phantom - 5 PCB coils energised sequentially
Dipole fit to localise and find inter-coil distance
Able to localise to within <1mm of error.
High spatial resolution and able to characterise cross-talk effects
Schofield et al., 2025 (in review)
X – CAD
X – Calibration
Holmes & Kennett, UoN
Tyler, Cerca Magnetics

## Slide 21: 384 channel Results – Visual Grating
Schofield et al., 2025 (in review)
See expected localisation and gamma band synchronisation during grating presentation
1.3x increase in amplitude with full array
Schofield,
Rier & Hill, UoN

## Slide 22: 384 channel Results – Spatial Leakage
Schofield,
Rier & Hill, UoN
Zero phase lag contributions between the two signals assumed to be leakage between the neighboring regions
Higher density should result in lower spatial leakage

## Slide 23: 384 channel Results – Spatial Leakage
Zero phase lag signal leakage shows >30% improvement with full array in some regions
Generally most improves areas with lower coverage and deeper regions
Schofield,
Rier & Hill, UoN
Schofield et al., 2025 (in review)

## Slide 24: Summary
High density OPM-MEG works!
Improves sensitivity
Improves spatial resolution
Accurate calibration is important
Higher channel counts maintain high density coverage

## Slide 25: Acknowledgements
Matt Brookes
Richard Bowtell
Mark Fromhold
Karen Mullinger
Matias Ison
Ryan Hill
Niall Holmes
Lukas Rier
Lauren Gascoyne
Jim Leggett
Jessikah Fildes
Dan Ferring
Dan Ford
Ewan Kennett
Gonzalo Reina Rivero
Ben Sanders
Clarise Castleman
Aditi Jain
Dominic Sims
Vishal Shah
James Osborne
Cody Doyle
David Bobela
David Woolger
Elena Boto
Eliot Dawson
Frank Worcester
Molly Rea
Holly Schofield
Zoe Tanner
Contact:
ppyjg12@nottingham.ac.uk
This work was supported by an Innovate UK biomedical catalyst grant (10037425) awarded to Cerca Magnetics Limited. We also acknowledge support from the UK Quantum Technology Hub in Sensing and Timing, funded by EPSRC (EP/T001046/1), the National Institutes of Health (Grant Number R01EB028772), an Innovate UK Collaborative R&D Award (Grant Number 104604), a Medical Research Council (MRC) Mid-Range Equipment grant (MC_PC_MR/X012263/1) and an Engineering and Physical Sciences Research Council (EPSRC) Healthcare Impact Partnership Grant (EP/V047264/1). Sensor development was made possible by funding from the National Institutes of Health (R44MH110288). We also acknowledge a Wellcome Collaborative Award in Science (203257/Z/16/Z and 203257/B/16/Z).