## Slide 1: Introduction to Optically Pumped Magnetometers (OPMs) For MEG
Peter D. D. Schwindt

## Slide 2: Outline
Introduction to MEG
SERF OPMs: What is SERF?
  Specific examples of SERF OPMs
How MEG measurements are made
  Sandia example
  Calibrating OPM arrays
  Coregistration
Other OPMs for MEG
Conclusion
  Comparison of MEG technologies

## Slide 3: Magnetoencephalography (MEG)
Lauri Parkkonen (Aalto University)
B. Maess, MPI for Human Cognitive and Brain Sciences
Postsynaptic currents flowing in the dendrites of the pyramidal neurons constitute the primary current ().
Both the primary and return currents contribute to the magnetic field sensed outside the subject’s skull.
Applications:
  Clinical: presurgical localization of epileptic foci
  Neuroscience: study normal and diseased brain, brain networks, etc.
Nerve Cell
Dendrites
Axon
Synapses
Journal of Neuroscience 2007, 27 (9) 2424-2432

## Slide 4: Functional Neural Imaging Technologies
Spatial
Resolution
Temporal
Resolution
EEG
MEG
fMRI
Poor
(~cm)
Great
(~ms)
Good
(mm to cm)
Great
(~ms)
Great
(~mm)
Poor
(~s)
fNIRS
Good
(mm to cm)
Poor depth
Poor
(~s)
Cost
$
$$$→$$
$$$
$$
Measurement
Paradigm
Neural
Activity
Neural
Activity
Blood
Flow
Blood
Flow
Naseer Noman, Hong Keum-Shik,”NIRS-based brain-computer interfaces: a review,” Frontiers in Human Neuroscience, 9 (2015)
Notes: Where MEG fits in with familiar neuroimaging techniques.
Problems with EEG – skull distortions.
Problems with fMRI – slow hemodynamic response, indirect measurement of brain activity

## Slide 5: Field strength
S. Seltzer Ph.D. thesis (2008)

## Slide 6: Traditional SQUID MEG vs. On-Scalp MEG
[1] Rainer Körber et al, “SQUIDs in biomagnetism: A roadmap towards improved healthcare,” in Superconductor Science and Technology 29(11):113001, DOI: 10.1088/0953-2048/29/11/113001
[2] Boto E, Holmes N, Leggett J, Roberts G, Shah V, Meyer SS, et al. Moving magnetoencephalography towards real-world applications with a wearable system. Nature. 2018;555(7698):657.
T ~ 180 C
~ 10 mm
[1]
[1]-modified
[2]
[2]
~ 15 mm
SQUIDs require liquid He (4 K).
Rigid helmet manufactured to fit 95% adult male subject’s head size.
Large sensor-source distance
  Diminished signal
  Less spatial complexity
On-scalp MEG enhances spatial resolution
Potentially wearable systems
  New research paradigms
Applications:
  Neuroscience research
  Brain Computer Interface (BCI)
  Clinical, e.g. epilepsy

## Slide 7: OPMs for MEG: Early work
Princeton/Twinleaf OPM MEG System
2006
NIST Miniature Magnetometer 2012
-now FieldLine Inc.
Sandia Atomic
Magnetometer
2010
QuSpin Inc. Magnetometer
2013

## Slide 8: Emerging Companies Using Zero-Field/SERF OPM
*
*Assuming this is a zero-field/SERF OPM

## Slide 9: Outline
Introduction to MEG
SERF OPMs: What is SERF?
  Specific examples of SERF OPMs
How MEG measurements are made
  Sandia example
  Calibrating OPM arrays
  Coregistration
Other OPMs for MEG
Conclusion
  Comparison of MEG technologies

## Slide 10: Atoms Used in OPMs

## Slide 11: Hyperfine Ground State
Ground State Hamiltonian:
  Spin angular momentum
  Nuclear angular momentum
  Total angular momentum
  , (87Rb), (133Cs)
  39K: MHz, , https://www.tobiastiecke.nl/archive/PotassiumProperties.pdf
  85Rb : GHz, , https://steck.us/alkalidata/rubidium85numbers.pdf
  87Rb : GHz, , https://steck.us/alkalidata/rubidium87numbers.1.6.pdf
  133Cs: GHz, , https://steck.us/alkalidata/cesiumnumbers.1.6.pdf
At low field the nuclear and electron spins are combined,
  ,
  ,
  2.8 MHz/G or 28 Hz/nT (gyromagnetic ratio)
  87Rb, 39K : 701 kHz/G or 7.01 Hz/nT
  85Rb: 467 kHz/G or 4.67 Hz/nT
  133Cs: 350 kHz/G or 3.5 Hz/nT

## Slide 12: OPMs: Three main physical effects
Atom-magnetic field interaction
  Zeeman effect
  Larmor precession
Atom-light interaction
  Optical pumping
  Faraday rotation
Decoherence effects
  Collisions
  Light
  Gradients

## Slide 13: The Rubidium-87 Ground State
Clock—measure hyperfine frequency
Magnetometer—measure Zeeman splitting
Gyromagnetic Ratio = 2p 7 Hz / nT
Atom Energy
mF
-2 -1 0 +1 +2
Hyperfine
splitting
Zeeman splitting
ћwHF
ћwL = 1/4mBB
F = 2
F = 1
F = I + S
I = 3/2, S = 1/2

## Slide 14: Optical Pumping
2P1/2
Lifetime ~ 27 ns
2S1/2
mS = -½
mS = ½
s+
2P1/2
2S1/2
mS = -½
mS = ½
s+

## Slide 15: Decoherence: Spin Exchange Relaxation Free (SERF)
T2 often limited by spin exchange collisions.
Near zero field and at high density, spin exchange rate >> spin precession frequency
  More F = 2 states, so on average atoms precess in F = 2 direction
  Precession rate is slowed down by nuclear spin and incomplete optical pumping
Record sensitivity: 160 aT / Hz1/2
  Atom shot noise: ~0.1 fT / Hz1/2
SERF in a ~6 mm3 scale cell: 5 fT / Hz1/2
H. D. Dang, A. C. Maloof, M. V. Romalis, Applied Physics Letters 97, 151110 (2010)
W. C. Griffith, S. Knappe, and J. Kitching, Optics Express, 18, 27167-27172 (2010).
F = 2
F = 1
mF
-2 -1 0 +1 +2
87Rb Ground States
Ensemble
Average
Single
Atom
Spin Exchange
Collision
Single
Atom
Ensemble
Average
Single
Atom
Spin Exchange
Collision
SERF Regime
High-Field Regime

## Slide 16: Transition between SERF and high-field regime
Cannot be modeled by the Bloch equation (without major modifications)
Density matrix analytical results (low polarization limit)
Where , is the linewidth (HWHM), is the Larmor frequency, is the bare atom Larmor angular frequency, and is the time between spin-exchange collisions.
Phys. Rev. A 71, 023405 (2005). DOI: https://doi.org/10.1103/PhysRevA.71.023405
Interesting read about the SERF regime: arXiv:2209.13086 [quant-ph]
sSE = 2 x 10-14 cm2
sSD = 9 x 10-18 cm2
Rubidium relaxation cross sections
Potassium vapor at 160, 180 and 200 °C
Twarge, CC BY-SA 4.0, https://creativecommons.org/licenses/by-sa/4.0, via Wikimedia Commons

## Slide 17: Optimizing sensitivity
1-spin exchange
2-add alkali spin destruction
3-add buffer gas spin destruction and diffusion
Kominis, Iannis K., Thomas W. Kornack, Joel C. Allred, and Michael V. Romalis. "A subfemtotesla multichannel atomic magnetometer." Nature 422, no. 6932 (2003): 596-599.
g = gyromagnetic ratio
= transverse coherence time
N = number of atoms
t = measurement time

## Slide 18: Density matrix for an atomic vapor-cell system
To carefully model the atomic spin dynamics, a density matrix treatment is required.
  Nuclear spin: and electron spin:
  Hyperfine interaction:
  Zeeman effect:
  Spin exchange:
  Spin destruction:
  Optical pumping:
  Diffusion:
  Phys. Rev. A 71, 023405 (2005). DOI: https://doi.org/10.1103/PhysRevA.71.023405
  Phys. Rev. A 58, 1412 (1998). DOI: https://doi.org/10.1103/PhysRevA.58.1412. Primary source

## Slide 19: Bloch Equations for an OPM
  Atoms precess under influence of magnetic field at rate given by the gyro magnetic ratio
  Electrons are oriented along circularly polarized pump beam (photon spin s) via optical pumping at rate
  Optical pumping causes depolarization of the collective magnetic moment
  Collisions cause atomic depolarization
  Probing the atoms causes depolarization of the collective magnetic moment at a rate
  Atoms diffuse through inert buffer gas (limits atom-wall depolarization)
By Lucas Vieira - Own work, Public Domain, https://commons.wikimedia.org/w/index.php?curid=1528090

## Slide 20: Bloch Equation for a zero-field SERF OPM
  Nuclear slowing-down factor depends on the polarization.
    Normally, However, in the SERF regime, both the gyromagnetic ratio and the relaxation rate depends on the atomic polarization.
  Note that (electron gyro magnetic ratio) is used
Scott Seltzer, “Developments in alkali-metal atomic magnetometry,” Thesis (Ph.D.), Princeton University, 2008.
Thomas Kornack, “A test of CPT and Lorentz symmetry using a potassium-helium-3 co-magnetometer,” Thesis (Ph.D.), Princeton University, 2005.

## Slide 21: Sensitivity Limits
Atom shot noise limit
Decoherence limits noise to 1/N1/2 scaling
Sensitivity improved by increasing T2 or N
Most AMs do not operate at this limit.
Photon shot noise limit
Most OPMs operate at or near the photon shot noise limit
Probe intensity contributes to T2 so Nph cannot be made arbitrarily high.
Magnetic
Sensitivity
g = gyromagnetic ratio
T2 = transverse coherence time
N = number of atoms
t = measurement time

## Slide 22: Zero-Field/SERF Magnetometer (Vector)
Alkali Vapor Cell
Randomly oriented atomic spins
Optical pumping
Spins align with the pump beam
Pump beam
Circular
polarization
Apply Small Magnetic Field
Spins precess due to magnetic field
B
Out of
plane
Detect with probe beam
Oriented spins rotate the polarization (Faraday rotation)
B
Input polarization
Output polarization
Probe beam

## Slide 23: Signals to Detect
Detect the Pump
Spin Polarization Bloch Equation
Bx = Bz = 0 Steady State Solution
B
Probe beam
Transverse Pump
B0+B1sin(wt)
Atomic Polarization, Pz
Pump Transmission
Atomic Polarization, Px
or Angle of Light Polarization
x
y
z

## Slide 24: Zero-Field SERF Magnetometer

## Slide 25: Outline
Introduction to MEG
SERF OPMs: What is SERF?
  Specific examples of SERF OPMs
How MEG measurements are made
  Sandia example
  Calibrating OPM arrays
  Coregistration
Other OPMs for MEG
Conclusion
  Comparison of MEG technologies

## Slide 26: Single beam SERF OPM
Magnetoencephalography with optically pumped magnetometers (OPM-MEG): the next generation of functional neuroimaging
Brookes, Matthew J. et al. Trends in Neurosciences, Volume 45, Issue 8, 621 - 634
Quarter Wave Plate

## Slide 27: Two-color pump/probe scheme
Optically pump on D1
Circularly polarized D1 pump beam
D1
795 nm
52S1/2
52P1/2
52P3/2
D2
780 nm
87Rb Fine Structure
Detect on D2
Linearly polarized 780 nm probe beam
Oriented spins rotate the polarization
Input polarization
Output polarization
Two optical resonances in Rubidium (fine structure)
  Use D1 for optical pumping and D2 for probing
  Based on: V. Shah and M. V. Romalis, PRA 80, 013416 (2009)
Dichroic optical components
To analyzer
D1 λ/4
D2 λ/2
Multiorder
Waveplate
780 nm
bandpass filter
Field modulation/phase sensitive detection
required to extract dispersive lineshape
To analyzer
D1 λ/4
D2 λ/2
780 nm
bandpass filter
B0+B1sin(wt)

## Slide 28: Magnetometer Signals
Blue trace: Magnetometer (lock-in) signal

## Slide 29: Linearity of OPMs
Bx
By
From Svenja Knappe

## Slide 30: Gen 2 Sensor Design
Other design changes:
  Reduced beam size: 2.5 mm FWHM to 2 mm
  Reduced nitrogen buffer gas: 600 Torr to 300 Torr
Result:
  Reduced temperature: 145 °C to 130 °C
  Reduced optical power:
    Pump: 38 mW to 4 mW
    Probe: 14 mW to 4 mW
l/2: 780 nm l/4: 795 nm
PM fiber for 780 & 795 nm laser
Vapor Cell
Detection Optics
Lens
DOE
18 mm
l/2
Filter:
Pass
780 nm
PBS
PBS
Photo-
diodes
Side View

## Slide 31: Gen 2 Oven design
Thin walled 3D printed oven walls
Aspen Aerogel Cryogel insulation: 5 mm
  15-25 mW/m/K (temperature dependent)
Distance from vapor cell center to external wall: 9 mm
Heating element:
  Co-fired ceramic heater
Demonstrated operating conditions:
  Power consumption: 6.1 W → ~3 W (all Cyrogel)
  Temp. underside: 178 °C
  Temp. stem: 151 °C
  Temp. outside: 47 °C → 37 °C when against back of hand
9 mm
Volume filled with foam except for beam paths
Insulation
Top co-fired ceramic heater

## Slide 32: Gen 2 Module Design
40 mm
266 mm

## Slide 33: On-Sensor coils for the Gen 2 Sensor
Designed by Joonas Iivanainen with bfieldtools
Coils:
  X: 2 coils
  Y: 2 coils
  Z: 4 coils
Channel
Locations

## Slide 34: Bandwidth
 | Ch 1 | Ch 2 | Ch 3 | Ch 4
3-dB BW (Hz) | 91 | 94 | 94 | 102

## Slide 35: Gen 2 Sensor Performance
 | Ch 1 | Ch 2 | Ch 3 | Ch 4
3-dB BW (Hz) | 92 | 88 | 98 | 87
Mag. Noise (fT/rt-Hz, 10-44 Hz) | 12.3 | 10.3 | 10.3 | 10.6
Grad. Noise (fT/rt-Hz, 10-44 Hz) | 6.9 | 6.3 | 5.4 | 5.3
Table Vibration

## Slide 36: MEG System: Block Diagram
24-ch, 6 sensor, OPM array
Human-sized magnetic shield
Custom laser source
Custom electronics
Custom LabView-based data acquisition and control software
Custom Matlab-based MEG data analysis pipeline

## Slide 37: MEG System: The 24-Channel Array
6-sensor, 24-channel array
Partially covers the left hemisphere

## Slide 38: MEG System: Signal Processing Pipeline
Read the Recorded MEG data (V)
Time Locked MEG Signal
MEG Signal Processing
Read the Recorded HPI Data (V)
Magnetic Resonance Image
Volume Conduction Model
ECD [3]
Source Location
INPUT
OUTPUT
ECD [3]
LCMV [4]
[1] R. Oostenveld, P. Fries, E. Maris, and J.-M. Schoffelen, "FieldTrip: Open Source Software for Advanced Analysis of MEG, EEG, and Invasive Electrophysiological Data," Computational Intelligence and Neuroscience, vol. 2011, p. 9, 2011.
[2] G. Nolte, "The magnetic lead field theorem in the quasistatic approximation and its use for magnetoencephalography forward calculation in realistic volume conductors," Physics in Medicine and Biology, Article vol. 48, no. 22, pp. 3637-3652, Nov 2003.
[3] Lutkenhoner B. "Dipole source localization by means of maximum likelihood estimation I. Theory and simulations,“ Electroencephalography and Clinical Neurophysiology, Volume 106, Issue 4, April 1998, Pages 314-321.
[4] B. D. Van Veen, W. van Drongelen, M. Yuchtman, and A. Suzuki, "Localization of brain electrical activity via linearly constrained minimum variance spatial filtering,“ IEEE Transactions on Biomedical Engineering, Article vol. 44, no. 9, pp. 867-880, Sep 1997.
A Combination of custom Matlab code and Fieldtrip toolbox [1]

## Slide 39: 20-Channels: 40 Measurements
Sequentially measure two transverse components
Calibrate each sensor to know its sensitive direction
Neighboring sensors interfere to change sensitive direction

## Slide 40: Summary of the method
A magnetic field sensor
  Gain (V/T)
  Orientation
  Position
A set of N magnetic field generators
Measure the sensor response to the ith generator
Model magnetic fields of the generators
Minimize the sum of squared errors between the measurements and the models to obtain , and
Sensor output:
Iivanainen, Joonas, et al. "Calibration and localization of optically pumped magnetometers using electromagnetic coils." Sensors 22.8 (2022): 3059.
Ryan M. Hill, et al.
‘Determining sensor geometry and gain in a wearable MEG system.” Imaging Neuroscience 2025; 3 imag_a_00535.

## Slide 41: MRI Co-registration and Forward Model
Four head localization coils close to the array
Polhemus Fastrak system to digitize
Relative sensor positions from CAD model
Head Localization Coils

## Slide 42: Optical coregistration
Zetter, R., Iivanainen, J. & Parkkonen, L. Optical Co-registration of MRI and On-scalp MEG. Sci Rep 9, 5490 (2019). https://doi.org/10.1038/s41598-019-41763-4

## Slide 43: Somatosensory/Auditory Evoked Magnetic Fields: Spatial Topographies
x
y
500 fT
x
y
500 fT
SEF
AEF

## Slide 44: OPM SEF/AEF Spatial Topography
N20m peak (~ 20 ms)
N100m peak (~ 100 ms)

## Slide 45: Source Localization: SEF
Red dot: Elekta Neuromag system and software. White dot: OPM data and Fieldtrip analysis software

## Slide 46: SERF OPMs: Formulating Gain/Sense-Angle Error
Cross Axis Projection Error (CAPE) Effects
Rotation of the sensitive axis
Reduction of gain
Phase errors
  Amir Borna, Joonas Iivanainen, Tony R. Carter, Jim McKay, Samu Taulu, Julia Stephen, Peter D.D. Schwindt, “Cross-Axis projection error in optically pumped magnetometers and its implication for magnetoencephalography systems,” NeuroImage, 247, 118818 (2022).

## Slide 47: Sensitive Axis Rotation: Simulation vs. Measurement
CAPE rotates the sensitive axis by 2.86 °/nT (measured) and 3.33 °/nT (simulated).
The slight difference between the two is attributed to different relaxation rates.

## Slide 48: Mitigation of CAPE
Zero the fields (easier)
  Zero the field at each sensor (< nT)
  Zero the field in the shield
Closed-loop, 3-axis OPM (harder)
  Upside: It’s a 3-axis measurement!
    Can we maintain good sensitivity?
  Upside: Fixed gain. Sensors are always calibrated.
  Downside: cross-talk between sensors.
QuSpin Gen 3 QZFM
Triax Variant: <23 fT/√Hz in 3-100 Hz band (typical 15 fT/√Hz all axis simultaneous)
Uses 2 perpendicular beams
Sandia Person-Sized Shield Coils
Nottingham Biplanar Coils
NeuroImage, 181, 760-774 (2018)
NeuroImage, 236, 118025 (2021)
NeuroImage, 252, 119027 (2022)

## Slide 49: Interference and crosstalk
Interference
  Leakage of the modulation field from one sensor to another
  Can easily be ameliorated through sensor calibration
Crosstalk in closed loop system
  Cancelling field at one sensor is sensed by another sensor
  Mitigation: Design feedback coils that minimize cross talk.
  Mitigation: A cross-talk matrix may need to be measured for each sensor configuration.

## Slide 50: Outline
Introduction to MEG
SERF OPMs: What is SERF?
  Specific examples of SERF OPMs
How MEG measurements are made
  Sandia example
  Calibrating OPM arrays
  Coregistration
Other OPMs for MEG
Conclusion
  Comparison of MEG technologies

## Slide 51: Other types of OPM
Unshielded MEG using scalar OPM
Helium-4 OPM
Sensitivity: 50 fT/rt-Hz
Closed-loop, triaxial operation
No cell heating: 10 mW RF discharge
Gradiometric Sensitivity: 15 fT/cm/rt-Hz
Measure field component parallel to ambient field
No crosstalk / magnetically silent
William Fourcault, et al., Opt. Express 29, 14467-14475 (2021)
Limes, M.E., et al., Phys. Rev. Appl. 14, 011002 (2020)

## Slide 52: Helium-4 OPM - principle
photo-detector
pump/probe
laser beam
Parametric resonances
along each direction
with 2 rf fields
and 1 laser beam
Signal (Arb. Units)
Magnetic field amplitude (nT)
0
50
100
-100
-50
-1
-2
-3
demodulation
F. Beato et al. Phys. Rev. A 98, 053431 (2018)
Linear polarization
Magnetic ALIGNED
state
6 cm
2x2 cm

## Slide 53: Pulsed Pump Magnetometer
Operate at nonzero background field
Extraordinary sensitivity without magnetic shielding
Free-induction decay frequency
PUMP PULSE
PUMP PULSE
High performance Optical Magnetic Gradiometers
Zero field SERF OPM
By
σ+ Light
87Rb Spin
θ∝B
Measure
spin angle
Measure
spin frequency
87Rb Spin
Probe
By
87Rb Spin
σ+ Pump
Pump
<1 µs
Probe
~1 ms
f = ɣB

## Slide 54: Outline
Introduction to MEG
SERF OPMs: What is SERF?
  Specific examples of SERF OPMs
How MEG measurements are made
  Sandia example
  Calibrating OPM arrays
  Coregistration
Other OPMs for MEG
Conclusion
  Comparison of MEG technologies

## Slide 55: SQUID vs OPM
SQUID
SERF OPM
Bx
By
Integrates over volume of laser beam
Sensitive axis determined by field modulation (Some SERF designs use laser beam geometry to define sensitivity axis.)
Integrates over the area of pickup loop
Sensitive axis determined by loop geometry
http://hyperphysics.phy-astr.gsu.edu/hbase/Solids/Squid.html

## Slide 56: SQUID vs OPM
 | Sensitivity | Bandwidth | Sensitive axis | Dynamic range (Max field)
Zero field, SERF OPM | 5-20 fT/rt-Hz | 80-150 Hz (open loop) 350 Hz (closed loop)** | 1, 2, or 3 axis (3rd axis from 2nd OPM) | ±2 nT (80 nT) open loop* ±150 nT (150 nT) closed loop**
He OPM | 50 fT/rt-Hz | 2 kHz (open loop) 1.35 kHz (closed loop) | 3 axis (3rd axis 6x worse sensitivity) | ±300 nT (unknown)
Scalar OPM gradio-meter—Unshielded | 15 fT/cm/ rt-Hz | 300 Samples/s (adjustable) | Parallel to am-bient bias field | Unknown (Earth’s field)
SQUID | 2-3 fT/rt-Hz | 1–5 kS/s (adjustable) | Normal to pickup loop | ±20 nT (Unshielded operation possible)
*Quspin Neuro-1: Neuro-1 flyer
**Fieldline: Alem O, et al.,Front. Neurosci. 17, 1190310 (2023)
Unknown = Unknown to this author

## Slide 57: Reporting results
Sensitivity results should be normalized by the bandwidth.

## Slide 58: Reporting results
Table Vibration
Gradiometrically inferred sensitivity = (Ch1 – Ch2)/
Gradiometer sensitivity = (Ch1 – Ch2)/1.8 cm
1.8 cm = gradiometer baseline

## Slide 59: Detection of epilepsy in children with OPMs
Odile Feys, Pierre Corvilain, Alec Aeby, Claudine Sculier, Niall Holmes, Matthew Brookes, Serge Goldman, Vincent Wens, and Xavier De Tiège, “On-Scalp Optically Pumped Magnetometers versus Cryogenic Magnetoencephalography for Diagnostic Evaluation of Epilepsy in School-aged Children,” Radiology 304:2, 429-434 (2022).

## Slide 60: Detection of epilepsy in children with OPMs

## Slide 61: Conclusions
OPMs are fundamentally different than SQUID sensors
  OPMs integrate field over a volume, not an area
  Bandwidth reduced compared to SQUID, at least for SERF OPMs
  Sensitivity of today’s OPM are reduced compared to SQUIDs
    Mitigated by being closer to the head
    Prospects for better OPM sensitivity: Best OPM 0.16 fT/rt-Hz
  OPM respond to all components of the field
    “Vectorization” methods needed to read out specific field components
    Three-axis detection possible
    Systematic errors arise from offset field
SERF and He OPM need a low-field environment
Scalar gradiometer operates in Earth’s field
OPMs are making a real world impact!
T ~ 180 C
~ 10 mm
[1]
[1]-modified
~ 15 mm

## Slide 62: Other references
Budker D, Jackson Kimball DF, eds. Optical Magnetometry. Cambridge University Press; (2013).
Labyt, Etienne, S. Tilmann, and Ronald Wakai. "Flexible High Performance Magnetic Field Sensors." Labyt, E., Sander, T., Wakai, R., Eds (2022).
Fabricant, Anne, Irina Novikova, and Georg Bison. "How to build a magnetometer with thermal atomic vapor: a tutorial." New Journal of Physics 25, no. 2 (2023): 025001.
Kimball, DF Jackson, E. B. Alexandrov, and D. Budker. "A Brief History of Optical Magnetometry." D. Budker (Cambridge University Press, Cambridge, UK, 2013) Chap 1: 1-10.

## Slide 63: Back up slides

## Slide 64: Properties of a Neutral Atoms
Quantized energy levels
p+
n = 1, l = 0
n = 2, l = 0,1
n = 3, l = 0,1,2
e-
Photon

## Slide 65: Properties of a Neutral Atoms
Spin and magnetic moment
  Electron spin
  Nuclear spin
p+
e-, ½
n = 1, l = 0
n = 2, l = 0,1
n = 3, l = 0,1,2
s = ½
s = -½
n = 1,
l = 0
m = -1
m = 1
F = 1
F = 0
m = 0
Zeeman effect
B
, ½

## Slide 66: Vapor cell optimization
Decoherence processes
  Diffusion (simplified to a lowest order rate)
    , cubic cell of length
    is the diffusion constant, and are the pressure and temperature at which is measured, is the buffer gas pressure at , and is the cell operating temperature
  Alkali-alkali spin destruction
    is the Rb density, is the Rb-Rb spin destruction cross section, and is the relative velocity
  Alkali-buffer-gas spin destruction
    is the buffer gas density, is the Rb-buffer-gas spin destruction cross section, is the relative velocity, and is the pressure at which the cross section is measured
  Noble gases typically have lower but excited state depolarization reduces optical pumping efficiency and allow radiation trapping
  Nitrogen has higher but advantageously quenches the excited state
= 5 mm
= 700 Torr
= 30 Torr
= 5 mm
= 30 Torr
= 150 C
Diffusion
Buffer gas
Spin Destruction
Alkali-Alkali
Spin Destruction
J. C. Allred, R. N. Lyman, T. W. Kornack, and Michael V. Romalis. "High-sensitivity atomic magnetometer unaffected by spin-exchange relaxation." Physical review letters 89, no. 13 (2002): 130801.

## Slide 67: Internal coils of the shield
Use the coils of the human-sized shield for calibration
Fluxgate measurement of the magnetic fields of the 18 shield coils
  Three-axis measurement at 108 positions around the sensor array
Model the magnetic fields using vector-spherical harmonics (VSH) expansion
  Maximum degree l = 5

## Slide 68: VSH model allows to calculate the currents to generate homogeneous and gradient fields

## Slide 69: Calibration algorithm
Get initial estimate on the sensor orientation , gain and position 0
  Use the 18 shield coils to excite homogeneous fields along x, y, and z
  Solve initial sensor orientation and gain by a matrix inversion
  Excite the five first-order gradients
  Given , solve initial sensor position 0 by a matrix inversion
Get final estimate on sensor orientation, gain and position
  Excite all 18 coils
  Solve the optimization problem using and 0 as the initial guess
    Use the full VSH models to take into account the whole spatial structure of the fields

## Slide 70: Validation of the method
Place three-axis fluxgate inside the sensor array
Use the algorithm to find the fluxgate parameters
Six positions inside the sensor array
  6×3 = 18 measurements (or sensors to calibrate)
The gains and relative positions/orientations of the fluxgate channels are known
  Gain (0.1 ±0.1%) V/µT
  Orthogonality ±0.5°
Sensys GmbH
Mean = 0.8%
Mean = 0.1°
Mean = 0.8 mm

## Slide 71: Simulation vs. Measurement, CAPE
OPM’s cross-axis projection error for a 25 Hz, sinusoidal signal with an amplitude of 0.33 nT.

## Slide 72: Impact of CAPE on MEG Signals: M20 Median Nerve Response
Wide bandwidth of SEF response at 20 ms (M20) reveals possible effects of CAPE.

## Slide 73: Impact of CAPE on OPM-MEG Systems’ Localization Capability
The mean and standard deviation (bars) for source localization error; the x-axis is the standard deviation of the remnant static magnetic field on the laser propagation axis selected from a normal gaussian distribution.
Poster Number: IM-29
Amir Borna