## Slide 1: Instrumentation for µSR
Kenji M. Kojima
TRIUMF & UBC: 2018~present
J-PARC: 2009~2018
U-Tokyo: 1998~2009
Columbia: 1993~1998 (visiting PhDPDF)
U-Tokyo: 1991-1993 (M.Sc)
Muon Spectroscopy school
14-21, 2026 Rigi-Kaltbad, Switzerland
https://indico.psi.ch/event/16947/

## Slide 2: Key Concepts/Takeaways:
Detection Fundamentals: Explanation of scintillators, photomultiplier tubes (PMTs), and how they convert positrons and muons into measurable signals.
Data Acquisition Chain: The role of discriminators, time-to-digital converters (TDCs), and how timing and dead-time considerations impact data quality.
Detector Arrays and Geometry: How multiple detector configurations (e.g., cylindrical arrays) are used to maximize solid angle coverage and handle high counting rates.

## Slide 3: Contents
Simplest µSR set up
Plastic Scintillator, PMT, Discriminator, Gate, TDC, DAQ cycle
Requirements to the detectors
  DC muon beam
  Pulsed muon beam
Fight against noise – what is the ground in circuits
Fight against temperature-based instability
Future capabilities in instruments
Summary

## Slide 4: Cosmic-ray muon life measurement
µ+ or µ-
Decay e+ or e-
Muon counter
Up counter
Down counter
Sample
PMT
Muon trigger
U counter e+/e- hit
D counter e+/e- hit
Plastic scintillator
t_0
t_hit (U)
t_hit(D)
Signal
timing
µ+  e+ + nm + ne
µ-  e- + nm + ne
NU(t) or ND(t)
t_hit (U or D) – t_0
tµ
= 2.2µs
*
Simplest µSR set up

## Slide 5: Plastic Scintillator
Anthracene dissolved in Acrylic plastic
Conduction band
LUMO
Valence band
HOMO
recombination
https://en.wikipedia.org/wiki/Anthracene
light
Charged particle
+
+
+
-
-
-
Example of
plastic scintillators
BicronLuxium solutions, Eljen, Torray, …
Fast, but poor energy resolution
due to light elements
https://luxiumsolutions.com/radiation-detection-scintillators/plastic-scintillators/bc400-bc404-bc408-bc412-bc416
Plastic scintillator

## Slide 6: Photo multiplier tube (PMT)
https://en.wikipedia.org/wiki/Photomultiplier_tube
Charged
Particle
(µ or e)
Anode: Electronic pulse signal
(Dynode: faster pulse signal)
High Voltage (~1-2.5 kV)
Current ~ 300-400 µA
Impedance
R=2.5kV / 0.4 mA
=6 MΩ
Typical gain ~ 106 (high)
Weak to magnetic fields
20ns
100mV
TM & V
V
TM

## Slide 7: Light guide and magnetic shield
https://cmms.triumf.ca/equip/lampf.html
https://www.hamamatsu.com/content/dam/hamamatsu-photonics/sites/documents/99_SALES_LIBRARY/etd/High_energy_PMT_TPMZ0003E.pdf
magnetic
shield
Light guide
~ 0.5m
TRIUMF LAMPF spectrometer
0.4T max  ~1mT at PMT

## Slide 8: SiPM (Silicon photo multiplier)
https://www.youtube.com/watch?v=eTPN7zceBa8
Japanese 5-transistor
radio ~1970s ?
https://en.wikipedia.org/wiki/Transistor_radio
https://www.hamamatsu.com/jp/en/product/optical-sensors/mppc.html
Hamamatsu MPPC
✅ Small & light
✅ Strong against magnetic field (~ Tesla)
✅ Low bias voltage (1-2.5kV  30-50V)
❌ Gain is temperature dependent
bias HV: -50mV/C
TRIUMF LAMPF spectrometer’s B-counter
20ns
5mV
400px 1mm sq,
bias=70.0V
for 69.82V
1st gen. MPPC S10362-11-050P
Sr90β（Emax=2.28MeV） on 1cm PS
1949 General Electric 115 AA5
Vacuum tube to solid state

## Slide 9: More about SiPM
Pixel array of Geiger mode APD
Laser on 400px SiPM HV=66.63 V
for 65.03V
2nd gen S12571-050P
https://hub.hamamatsu.com/us/en/technical-notes/mppc-sipms/a-technical-guide-to-silicon-photomutlipliers-MPPC-overview.html
Impedance: 65 V / ~4µA =16 MΩ
M.V. de Toro Sánchez et al 2023 J. Phys.: Conf. Ser. 2462 012014
N
P
light
-HV
GND
++++++++
- - - - - - - - -
- +
avalanche
avalanche

## Slide 10: Key Concepts/Takeaways:
Detection Fundamentals: Explanation of scintillators, photomultiplier tubes (PMTs), and how they convert positrons and muons into measurable signals.
Data Acquisition Chain: The role of discriminators, time-to-digital converters (TDCs), and how timing and dead-time considerations impact data quality.
Detector Arrays and Geometry: How multiple detector configurations (e.g., cylindrical arrays) are used to maximize solid angle coverage and handle high counting rates.

## Slide 11: Discriminators (Analog signal to Digital timing pulse)
µ+ or µ-
Decay e+ or e-
TM counter
Up counter
Down counter
Sample
PMT
Muon trigger
U counter e+/e- hit
D counter e+/e- hit
Plastic scintillator
t_0
t_hit (U)
t_hit(D)
Signal
timings
Analog
Digital
Discriminator
D E Pooley et al 2023 J. Phys.: Conf. Ser. 2462 012028.
Peak detect ( super-MuSR@ISIS)
Misses signals in high rate
Lecroy 623
discriminator

## Slide 12: Time to Digital Converter (TDC)
NU(t) or ND(t)
t_hit (U or D) – t_0
tµ
= 2.2µs
*
Analog
Digital
µ+ or µ-
Decay e+ or e-
TM counter
Up counter
Down counter
Sample
PMT
Muon trigger
U counter e+/e- hit
D counter e+/e- hit
Plastic scintillator
t_0
t_hit (U)
t_hit(D)
Signal
timings
Discriminator
Start / Stop
Reset
 TDC start
 TDC stops
Where is the reset signal ?
When do you read data?
https://en.wikipedia.org/wiki/Stopwatch
https://hif.wikipedia.org/wiki/Computer
Hit data to
DAQ-computer
for analysis

## Slide 13: Gate signal, pile up and DAQ cycle
Synchrotron (Pulsed muons)
PSI
TRIUMF
Cyclotrons (DC muons)
ISIS
J-PARC RCS
Muon with random arrival times
Limit ~40k muons/second
muons
time
Gate = Time window ~10µs
time
 20 ms (ISIS) to 40 ms (J-PARC)
TDC reset
Thousands of muons / frame
Record times of multiple e+ hits
When µ+ dies away (~32-65µs) TDC reset
TDC start
TDC stop
if there is a e+ hit
muons
1st Muon Community Meeting (2021): https://indico.cern.ch/event/1030726/

## Slide 14: Cosmic-ray setup ≒ DC-µSR Counter geometry
https://www.psi.ch/sites/default/files/import/smus/InstrumentGpsEN/User_Guide_present_version.pdf
µ+
e+
Muon
Counter
Up counter
Down + veto counter
Sample
PMT
Plastic scintillator
t_0
t_hit (U)
t_hit(D)
Signal
timings
µ+
e+
PSI-GPS counter geometry
Count rate consideration
Muon counter: 40k µ+ / sec
Average 1/40k = 25µs separation
1 e+ / 1 µ+ = 40k e+/sec at most.
TRIUMF LAMPF electronics with 10 counters. NIM electronics is OK.

## Slide 15: Pulsed µSR: no µ+ counter, but many e+ counters
Count rate consideration
ISIS ARGUS 40M e+/h (/3600/50Hz) ~220 e+ /frame
J-PARC S1/D1 200M e+/h (/3600/25Hz) ~ 2200 e+ / frame
These e+ are generated in ~ tµ=2.2µs:
Instantaneous rate 2200 e+ / 2.2µs = 1G counts / sec !
Fragments: 100 counters  10 Mcps, or 22 e+ / frame
1000 counters  1 Mcps or 2.2 e+ / frame
192 ch
64 ch
64 ch
ARTEMIS (J-PARC MLF S1)
D1- spectrometer (MLF D1)
640 ch pair
10 counters  2 racks (TRIUMF)
64 counters  12 racks (?)
1280 counters (640 pairs)  128 racks (!)
J-PARC MLF spectrometers uses ASIC-ASD + FPGA
SiPM
32 ch of 10x10x12mm
BC408 scintillator + Y11 fibre
on a Hamamatsu MPPC sensor
TDC + memory
ASD (Amp-Shaper-Discriminator)
Kalliope
K M Kojima et al
J. Phys.: Conf. Ser. 551 012063 (2014 )

## Slide 16: Electronics response required for DC vs. Pulsed µSR
Electronics specs | DC | Pulsed
Count rate | 40kcps, continuous a few counters ~1e+/10µs | 1 Gcps, instantaneous 100-1000s counters ~10e+/frame/counter
Rise-time resolution | very important | not important
Tail recovery time | any | as short as possible
Base-line stability | any | must be flat, no base-line shift
Desired analog pulse |  | 
Long tail is OK
Undershoot is OK
Fast tail
NO undershooting
Fast rise time
No jitter
Slow rise
is OK
CFD (Constant Fraction Discriminator)

## Slide 17: Constant Fraction Discriminator (CFD) – DC-µSR
Fixed threshold discriminator
Depending on the pulse height, the leading-edge timing changes
Time walk
Principle of CFD
https://en.wikipedia.org/wiki/Constant_fraction_discriminator
Timing of the zero-crossing point is independent to the pulse height

## Slide 18: Pile up rejection (DC) vs. correction (Pulsed)
Gate = Time window ~10µs
Count loss in high rate
The lost counts must be corrected, assuming a deadtime model.
Or go to peak detect scheme
If a 2nd muon comes in within the gate window, such event must be discarded.
If more than one e+ hit is recorded within the gate window, such event must be discarded.
Fixed deadtime t model and correction works for Volume2014 amplifier
Ncorr = Nobs/(1-Nobs t)
KMKojima et al. JPS Conf. Proc. , 011062 (2018)
D E Pooley et al
2023 J. Phys.: Conf. Ser.
B counter
8000
Time (ns)
TRIUMF Dilution Refrigerator
Nbg
~ 3% N(t0)
N(t0)
Background level goes up or has time dependence, if wrong.

## Slide 19: Counter geometry, beamline and cryostat etc.
s
Small counters: large asy, small rate
Large counters: small asy, large rate
Figure of merit ~ A2S
Beam duct for muon beam
Flypast chamber (pulse-µSR)
LF magnet coils
Cryostats
High LFs focuses the beam and
alter the e+ trajectory  alpha-shift
Full simulation of beam optics and counter geometry
 Peter Baker’s Super-MuSR talk on last day

## Slide 20: Contents
Simplest µSR set up
Plastic Scintillator, PMT, Discriminator, Gate, TDC, DAQ cycle
Requirements to the detectors
  DC muon beam
  Pulsed muon beam
Fight against noise – what is the ground in circuits
Fight against temperature-based instability
Future capabilities in instruments
Summary

## Slide 21: How do you define the ground of circuits ?
Where is
the ground of
my circuits??
Worst BL component in noise: Magnetic Kicker
400 A 50 kV
=2000 kW pulse
SiPM
0 4000

## Slide 22: Noise propagation
LC connected to small PS  noise level 1/10
Detector connected to the BL ground  1/350
Definition of the ground:
the largest conductor of the system
= LC (longitudinal Coil)
Can we use LC as the ground ?
(could not just connected by wire,
because LC is energized to 1000A 100V)
~1/350
~1/700

## Slide 23
How to engage LC as the ground
L~10mH, C=100nF: Resonance f=5kHz,
ie 200µs period. Not visible to µSR (0~30µs)
10,000
0
3000
1
Capacitor is transparent in high frequency
Impedance Z=1/iωC
LC is capacitance-coupled to the ground, where the noise is bypassed via C.

## Slide 24: Fighting against the T-dependent SiPM gain
When the LF coil (1000A x 100V=100kW) is energized, the spectrum start to drift, due to the thermal coefficient of SiPMs.
Coil surface: +10 ℃
SiPM: +5~10 ℃
-50mV/ ℃ x 5 = -0.25V bias change
0.3
0.1

## Slide 25: A copper can with cooling water pipes are installed inside and outside of the detector modules. With a fixed T chiller ($2k), the temperature rise was reduced to +2℃ from +10℃ of no active cooling. With a PID controlled chiller ($20k), the temperature is within 0.5K.
PID controlled chiller
No apparent drift observed
0.2
0.15
0.2
0.15

## Slide 26: Future capabilities
Muons counter hits
Gate = Time window ~10µs
TDC reset
TDC start
TDC stop
if there is a e+ hit
(1) Pileup limitation of 40k muons / sec.
What if Muon counters are pixeled, and
the source muon is identified for each positron?
40k muons / sec limit is lifted
Muon spin imaging will become possible
e+
e+
Zaher Salman’s talk
on vertex reconstructed
µSR at PSI
(2) Maximizing the efficiency of pulsed µSR
Peter Baker’s talk
on super-MuSR at ISIS
Peak detect
Discriminator
Not just minimizing the count loss, but also maximizing the beam intensity, and optimizing the counter geometry, etc
…

## Slide 27: Summary (µSR instrumentation)
Introduced:
Plastic Scintillator, PMT, Discriminator, Gate, TDC, DAQ cycle
Difference for DC and pulsed muon sources
Noise mitigation, SiPM thermal compensation
Most of the problems are already solved before the user comes.
Enjoy science based on the instruments !

## Slide 28: Back ups

## Slide 29: Time over threshold (TOT)
Fixed threshold discriminator
The duration of the digital pulse reflects the pulse height, if the pulse exceeds the threshold.
Most of the NIM discriminators
(eg. LeCroy 623) has fixed widths.
Not providing TOT.
Convenient method to estimate the pulse height

## Slide 30: Constant Fraction Discriminator (CFD) – DC-µSR
Fixed threshold discriminator
Depending on the pulse height, the leading-edge timing changes
Time walk
Principle of CFD
https://en.wikipedia.org/wiki/Constant_fraction_discriminator
Timing of the zero-crossing point is independent to the pulse height

## Slide 31: Characteristics pixel size dependence
R１
C１
R
C
Bias (-Vcc)
Vout
GND
R1=10k, C1=100nF
R=10k, C=100nF
20ns
5mV
Noise filter
100px, bias=70.6V
for 70.17V
400px, bias=70.0V
for 69.82V
1600px, bias=73.5V
for 71.74V
100µm pitch
50µm pitch
25µm pitch