## Slide 1: Muons at extreme conditions: pressure
Rustem Khasanov
Rigi Kaltbad, 16 January 2025

## Slide 2: Common tuning parameters
T
B
P
P-T phase diagram of water
Room temperature superconductivity
Wikipedia

## Slide 3: Principle of mSR under pressure experiments
Muon momentum tuning
Beam-width tuning
Courtesy of Daniel Andreica

## Slide 4: mSR under pressure: basic principles
Khasanov JAP, 132, 190903 (2022)

## Slide 5
mSR under pressure
1. Muon beam-line – fast muons with tunable energy
2. mSR Spectrometer
3. mSR pressure cells

## Slide 6
mSR under pressure
1. Muon beam-line – fast muons with tunable energy
2. mSR Spectrometer
3. mSR pressure cells

## Slide 7: Muon beam

## Slide 8: Decay muon beam-line

## Slide 9: mE1 decay beam-line at PSI
The use of the spllited Quadrupolar magnet (QSK81) allows to collect muons with turned spins. This a unique possibility which is accessible for decay muon beam-lines only.
Momentum acceptance (FWHM) 3%
Pion momentum range [MeV/c] 200--125
Muon momentum range [MeV/c] 125– 60
Rate of positive muon [mA-1s-1] 3x107—6x107
Spot size (FWHM) 39 x 28 mm

## Slide 10: Asymmetrically driven quadrupole
FL
FL
FL
FL

## Slide 11
mSR under pressure
1. Muon beam-line – fast muons with tunable energy
2. mSR Spectrometer
3. mSR pressure cells

## Slide 12: Muon spectrometer (GPD)
Individual Detector
Detector head
Detector setup
GPD spectrometer

## Slide 13
mSR under pressure
1. Muon beam-line – fast muons with tunable energy
2. mSR Spectrometer
3. mSR pressure cells

## Slide 14: Pressure cell construction: compound cylinder
Single wall cell
Double wall cell
Three wall cell
a
b
c
c1
c2
For a=6 mm and b=24 mm, pmaxs ÷ pmaxd ÷ pmaxt = 1 / 1.6/ 1.96

## Slide 15: Construction material suitable for mSR
 | CuBe | TiAl6V4 | NiCrAl | MP35N
Yield strength | 1.1 Gpa (300 K) | 1.05 Gpa (300 K) | 2.06 Gpa (300 K) | 2.15 GPa (300 K)
Young modulus | 131 GPa (300 K) | 97 Gpa (300 K) | 190 Gpa (300 K) | 215 Gpa (300 K)
 | WC | cBN | SiC | ZrO2-Y2O3 | Al2O3-ZrO2 | Si3N4
Compressive strength | 5.0-11.0 Gpa | 2.9 GPa | 7.6-8.3 GPa | 2.20 GPa | 4.7 GPa | 5.1-5.5 GPa
Young modulus | 600-670 Gpa |  | 918 GPa | 210 Gpa | 357 GPa | 241 GPa
Nonmagnetic Alloys
Sintered materials
Strong enough to hold the pressure
Should not have “strong” mSR response
Should have temperature independent response

## Slide 16: Three-wall pressure cell construction
Top fixation bolt
Sample volume
Bottom fixation bolt
Pressure seals
pmax(RT) ~3.3 Gpa, pmax (LT) ~ 3.0 GPa

## Slide 17: Pressure determination, pressure probes
Contact (feedthroughs)
Contactless
Resistivity, AC susceptibility, NMR, NQR,
specific heat, optical …
Optical, AC susceptibility, NMR, NQR, specific heat,
Neutron scattering (equation of state)…
Substantial part of the pressure cell volume is occupied by the pressure indicator

## Slide 18: Double volume piston-cylinder cell
Optical window
Piston
Teflon mushroom
Optical volume
Sample volume

## Slide 19
USB
520 nm pigtailed laser (PLM520.0MMF01)
Optical spectrometer (Ocean Optics HR400 or HR4PRO)
Control PC
Double‐volume pressure cell

## Slide 20: Laser safety

## Slide 21: Safety concern: Laser vs. LED light
cooling fan
heatsink
LED
fiber plug
(SMA905)
lens
(a)
(b)
(c)
lens holder
~1.5mm

## Slide 22: LED light sources for fluorescense

## Slide 23: Ruby and Sr tetraborite

## Slide 24: Diamond anvil setup
LED
source
Optical
spectrometer

## Slide 25
Uniaxial pressure experiments

## Slide 26: Uniaxial measurement stage
Ghosh et al., RSI 91, 103902 (2020)

## Slide 27: Schematic of uniaxial pressure cell
Ghosh et al., RSI 91, 103902 (2020)
Sample holder: detachable
Maximum force: up to 1000N
Sample dimensions: down to 3x3x0.2mm3
T-range: 0.27K – 300K

## Slide 28
Scientific examples
1. The uniaxial and hydrostatic pressure effects on TRSB in Sr2RuO4
2. Pressure-induced critical and multicritical points in frustrated spin liquid
3. Pressure split of the density wave transition in RP nickelate La3Ni2O7

## Slide 29
Scientific examples
1. The uniaxial and hydrostatic pressure effects on TRSB in Sr2RuO4
2. Pressure-induced critical and multicritical points in frustrated spin liquid
3. Pressure split of the density wave transition in RP nickelate La3Ni2O7

## Slide 30: Time-reversal symmetry breaking in Sr2RuO4

## Slide 31: Uniaxial strain
Grinenko et al., Nature Phys. 2021

## Slide 32: Hydrostatic pressure experiments
Grinenko et al., Nature Comm. 2021

## Slide 33: Combined graph data

## Slide 34
Scientific examples
1. The uniaxial and hydrostatic pressure effects on TRSB in Sr2RuO4
2. Pressure-induced critical and multicritical points in frustrated spin liquid
3. Pressure split of the density wave transition in RP nickelate La3Ni2O7

## Slide 35: (C4H12N2)Cu2Cl6 (PHCC)

## Slide 36: p-T phase diagram of PHCC

## Slide 37
Scientific examples
1. The uniaxial and hydrostatic pressure effects on TRSB in Sr2RuO4
2. Pressure-induced critical and multicritical points in frustrated spin liquid
3. Pressure split of the density wave transition in RP nickelate La3Ni2O7

## Slide 38: Superconductivity in La3Ni2O7 – July 2023
H. Sun et al., Nature 621, 493 (2023)

## Slide 39: ZF-mSR, ambient pressure
The response is typical for magnetically ordered sample
Two contributions – the fast and the slow oscillating – are present
Confirmation of the results of Lei Shu group [K.W. Chen et al., arxiv:2311.15717] with, however, some differences.

## Slide 40: Commensurate vs. Incommensurate SDW order
The commensurate magnetic order is characterized by a well defined internal field value at the muon stopping site(s). The magnetic field distribution is, typically, consists of a peak(s) of Gaussian or Lorenzian shape
The incommensurate magnetic order is characterized by a zeroth-order Bessel function, or a ‘double-peak’ type of distributions
A. Amato and E. Morenzoni, Springer Verlag 2024

## Slide 41: mSR experiments under pressure

## Slide 42: La3Ni2O7 phase diagram
The DW-like anomaly in resistivity is progressively suppressed with increasing pressure, which may suggest competition with the superconducting order.
The SDW transition temperature increases with increasing pressure
The distinctive pressure-induced response in La3Ni2O7 implies the simultaneous existence of two separate orders in its ground state: magnetic and most likely charge density wave orders.
In cuprates or hole-doped nickelates like La2-xSrxNiO4, the relationship between spin and charge orders is notably intricate, with both orders being strongly intertwined and exhibiting similar responses to external parameters.

## Slide 43: Magnetic structure(s) of La3Ni2O7
Neutron scattering: T. Xie arXiv:2401.12635
RIXS: X. Chen et al., arxiv:2401.12657
NMR: Z. Dan et al., arxiv:240203952

## Slide 44: Muon sites and magnetic structures
La
Ni
O
µ+
Double-stripe (DS)
DS with vacancies 1
DS with vacancies 2
The DFT+m calculations reveal the presence of one
muon stopping site at (0.495; 0.006; 0.170).

## Slide 45: Field distributions and the ordered moment value
(b)
(c)
(d)
(e)
The dipole-field simulations suggests that some of the Ni sites must carry no magnetic moment, otherwise the slow precession component (corresponding to a low magnetic field at the muon site) is not observed.
La3Ni2O7 must host a decoupled charge- and spin-density wave, analogous to a stripe-type order in cuprate high-temperature superconductors.
By considering all possible configurations with missing Ni magnetic moments, an upper limit of the moment on the Ni site would stay mNi = 0.59(5) mB per f.u. for DSv1 or, mNi = 0.29(5) mB per f.u. for DSv2

## Slide 46
My thanks go to
PSI: Matthias Elender, Zurab Guguchia, Markus Bendele, Alexander Maisuradze, Zurab Shermadini, Ritu Gupta, Stefan Hollenstein, Gediminas Simutis, Debarchan Das, Andrea Raselli, Katja Pomjakushina, Dariusz Gawryluk, Pavel Naumov, Marek Bartkowiak, Nicola Casati, Thomas Hicken, Jonas Krieger, Elvezio Morenzoni, Hubertus Luetkens, Alex Amato …..
Uni Sorbonne: Stefan Klotz
Uni Edinburgh: Konstantin Kamenev and his lab
Ill: Eddy Lelievre-Berna and his group
ISIS: Oleg Kirichek and his group
Ames Lab: Sergey Bud’ko, Paul Canfield
Institute for High Pressure: Ravil Sadykov