## Slide 1: Testing the new Low Energy Germanium Detectors (LEGe)

## Slide 2: Setting up/ Parameters
Set the gain of the Intelligent Preamplifier at its maximum option: x10
Trapezoid filter Parameters
LEGe 1 |  |  | 
Trapezoid Parameters | Tau (ns) | Gap | Peaking
Used | 49.19 | 300 | 2600
Manufacturer |  | 200 | 1800
LEGe 2 |  |  | 
Trapezoid Parameters | Tau (ns) | Gap | Peaking
Used | 48.96 | 250 | 2100
Manufacturer |  | 200 | 1800

## Slide 3: Tau Parameter
Taking long traces with 25000 buffer length

## Slide 4: Gap and Peaking
Running the trapezoid filter with Michaels Trapezoidal Optimization code.

## Slide 5: Gap and Peaking
Choosing gap at 300 for LEGe1 and 250 for LEGE2.

## Slide 6: Gap and Peaking
Fixing the Gap we study the minimum resolution for the Peak parameter

## Slide 7: Spectra
Using three sources at the same time: Fe-55, Co-57 and Am-241

## Slide 8: Fitting the Peaks
x 40.8

## Slide 9: Resolution
LEGe 1 |  |  | 
Isotopes | Fe-55 | Am-241 | Co-57
Energy (keV) | 5.9 | 59.54 | 122
Calculated (eV) | 235 | 377 | 517
Manufacturer (eV) | 220 | - | 525
LEGe 2 |  |  | 
Isotopes | Fe-55 | Am-241 | Co-57
Energy (keV) | 5.9 | 59.54 | 122
Calculated (eV) | 235 | 400 | 542
Manufacturer (eV) | 220 | - | 525

## Slide 10: Resolution

## Slide 11: Efficiency
Calculated intrinsic efficiency, considering the X-ray attenuation for Be window and plastic support around sources 1mm.

## Slide 12: Efficiency
Calculated intrinsic efficiency, considering the X-ray attenuation for Be window and plastic support around sources 1mm.

## Slide 13: Discussion Topic: The peaking normalization
When we removed at the BankDecoder.cpp the peaking normalization from trigger.energy we got a much better resolution, especially for Fe-55 source

## Slide 14: Discussion Topic: The peaking normalization
When we removed at the BankDecoder.cpp the peaking normalization from trigger.energy we got a much better resolution, especially for Fe-55 source
No peaking normalization: FWHM(eV): 0.228
Peaking normalization: FWHM(eV): 0.314