Machine Report Week 08/2025
Thomas Schietinger
SwissFEL Exchange Meeting, 24 February 2025

---
Calendar(s)

---
Photon (pulse) energies
 ● Photon energy (keV)
 ● Photon pulse energy (µJ)
Mon, 17-Feb, 08:00 – Mon, 24-Feb, 08:00
Aramis
7.1 keV for Alvra, ~1.2 mJ
12 keV for Cristallina user experiment, ~800 µJ
Setup
short-pulse mode
Athos
Machine development
780 eV (LH/LV) for Furka user experiment
~ 1.4 mJ
Setup
680 eV/510 eV (LH) for Maloja user experiment
~ 1.8 mJ @ 680 eV

---
Aramis short pulses
Ph. Dijkstal
- Nominal beam 7.1 keV, 1.1 mJ, 42 fs (fwhm).
- Decompression of electron beam for narrower bandwidth.
- Pulse shortening with vertical beam tilting from dechirper (horizontal plates).
- Measurement of short pulse (refined analysis) gives ~17±3 fs.
comm. ELOG 59662
comm.
ELOG 59665

---
Athos short pulse
Ph. Dijkstal
- Nominal beam 680 eV (LH) with 1.8 mJ (after optimization). Photon pulse length 36 fs (fwhm).
- Pulse shortening with vertical beam tilting from skew quad in dogleg.
- Large horizontal tilt in beam prevents good measurement with transverse deflector.
- Only upper limit of 25 fs (fwhm) can be given.
comm. ELOG 59670

---
Issues last week
Causing machine downtime:
- AC synchronization error (Monday during setup, Tuesday 21:30):
  knocks out all RF stations. RF stations come back, but synchronization
  error persists. To be fixed this week.
- Repeated gun interlocks (Friday, 10:50, every ten minutes):
  reason is waveguide pressure interlock, temporarily fixed by lowering
  alarm threshold. RF team will refill SF6 during Easter shutdown.
- Instability in Mizar laser (Friday 13:30): intervention by Carlo shows that
  the laser cavity is slightly misaligned, leading to unstable output. Compensated by increase in diode current. Realignment of laser cavity should be scheduled soon.
Other issues (no significant downtime)
- Soft MPS alarms due to ICT calibration error (Tuesday, 15:00): reload of calibration constants is needed.
- Alcor LAM failure (Wednesday, 13:15): laser pikett fixes it
- Mizar LAM failure (Thursday, 19:30): needs power cycle and manual reset of motor position (done by operators)