## Slide 1: Vertex Calibration + Pixel Tuning
Mu3e collaboration meeting
Zurich
Thomas Senger

## Slide 2: Calibration
IV scan
VDAC power consumption
Signal transmission
BLPix vs VNOutPix (Eficiency and Stability)
VPTimerDel & VNTimerDel

## Slide 3: IV Scan
IV curves performed after extraction
  5 partitions are working
  2 have no current, one in HV compliance
Factor of ~1.5 difference in HV current at 0°C vs 20°C
Temp Diode difference max 9°C

## Slide 4: IV Scan
IV curves performed after extraction
  5 partitions are working
  2 have no current, one in HV compliance
Factor of ~1.5 difference in HV current at 0°C vs 20°C
Temp Diode difference max 9°C

## Slide 5: DAC Power Consumption VPDAC
Each line represents one Mupix sensor
Would expect 2mA per DAC
Observe more or less flat distribution
Resolution of DCDC boards to low to see differences

## Slide 6: DAC Power Consumption BLPix
Each line represents one Mupix sensor
Would expect in total 80mA increase per chip
Observe more or less flat distribution
Resolution of DCDC boards to low to see differences

## Slide 7: DAC Power Consumption ref_vss
Each line represents one Mupix sensor
Expect total increase per chip O(150mA)
Data is compatible with QC data
Operation at 169  in plateau region

## Slide 8: Signal transmission scan 8b10b errors

## Slide 9: Signal transmission scan disparity errors

## Slide 10: Signal transmission scan optimal settings
Eye diagramm measurement at UZH
Results of the Vertex Detector

## Slide 11: BLPix vs VNOutPix 8b10b errors

## Slide 12: BLPix vs VNOutPix Active pixels

## Slide 13: BLPix vs VNOutPix Active pixels (with 8b10b error exclusion)

## Slide 14: Optimal Parameters
Default settings:
  VNOutPix=5
  BLPix=80
Optimal settings found at
  No 8b10b errors
  Most active pixels

## Slide 15: BLPix vs VNOutPix Current consumption + Temperature

## Slide 16: VPTimerDel
Three chips as example
Want to use full ToT range  can be controlled through VPTimerDel
Default Value was 10 before the beamtime
  Ckdivend2 put 1531
Low VPTimerDel  full ToT spectrum
  Too low: Voltage drop visible, pixels are not readout (Current Interpretation)
Set cut to ToT=30 + full hitmap

## Slide 17: VNTimerDel
Three chips as example
No difference in ToT for different VNTimerDel

## Slide 18: Tuning

## Slide 19: Reminder: VPDAC & TDAC
Small threshold plateau region of high efficiency
Goal is low threshold operation
  Noise at the edges is dominant
  Chips get unstable at low threshold
VPDAC determines the tuning current range
increases power consumption of the chip
8 TDAC bits
  0_1masking_3ThLow_3ThHigh

## Slide 20: Reminder: VPDAC & TDAC
Small threshold plateau region of high efficiency
Goal is low threshold operation
  Noise at the edges is dominant
  Chips get unstable at low threshold
VPDAC determines the tuning current range
increases power consumption of the chip
8 TDAC bits
  0_1masking_3ThLow_3ThHigh
Has to be determined

## Slide 21: Algorithm: Brute force tuning

## Slide 22: Noise scan

## Slide 23: Noise scan
Noise only appears below ThHigh=130
  Was 160 in QC stand
Only shown here from ThHigh = 126
Noise first visible at the edges
Chips seem to behave similar on same ladder
8b10b flooded on certain chips at low threshold
At 117 metal structures are visible
  Also visible during QC

## Slide 24: Noise scan
Noise only appears below ThHigh=130
  Was 160 in QC stand
Only shown here from ThHigh = 126
Noise first visible at the edges
Chips seem to behave similar on same ladder
8b10b flooded on certain chips at low threshold
At 117 metal structures are visible
  Also visible during QC

## Slide 25: Find Optimal VPDAC

## Slide 26: VPDAC scan

## Slide 27: Optimal VPDAC values

## Slide 28: Alpha estimation
Only rough alpha estimation
  comparison of noise scan and optimal VPDAC estimation
Need to mask edges and perform noise scans at different TDAC values at fixed VPDAC
  As shown in Wengen
TDAC writing takes currently too long
  50ms per writing
  Does not work for 10ms
  Can this be solved?

## Slide 29: Alpha estimation
Only rough alpha estimation
  comparison of noise scan and optimal VPDAC estimation
Need to mask edges and perform noise scans at different TDAC values at fixed VPDAC
  As shown in Wengen
TDAC writing takes currently too long
  50ms per writing
  Does not work for 10ms
  Can this be solved?

## Slide 30: Algorithm: Not so Brute force tuning Part 1

## Slide 31: Algorithm: Not so Brute force tuning Part 2

## Slide 32: Next steps and open questions
Perform noise scan at different VPDAC values at different TDAC values with masked edges
  Determine alpha per chip – are there chip to chip differences
Perform full tuning with optimal TDAC/VPDAC values
How to not be influenced by 8b10b errors
  More steps seem not to help
What is our goal threshold
  Do we also tune the metal layers (ThHigh=117)
Reduce time of TDAC writing (currently >90s)

## Slide 33: Weird Mupix structures

## Slide 34: Weird Mupix structures
QC