## Slide 1: Advanced finite element modeling to support the experimental test MAG-S.01.04-T023-D001
F. Giorgetti
WPMAG 2025 Final Meeting, 15-Jan-2026

## Slide 2: Outline
Overview of the ENEA PC-RINGS Testing Facility.
FE model to support the experimental test and determine the testing cycle amplitude.
Code developed for crack growth analysis.

## Slide 3: ENEA PC-RINGS Testing Facility
Facility Key features
Based on eighteen (18) hydraulic cylinders with maximum pulling capability of 57tons and a maximum stroke of 70mm each, thus total σHmax is 1.6MN.
Position control precision is 0.1mm (below 0.1mm radial movement accuracy cannot be guaranteed).
Mounting and demounting procedures are very time-consuming (intermediate disassembly not recommended).
Cycle duration 10s÷20s (after software modification)
Testing requirements
Ring mock-up dimensions: Internal diameter of 1005mm±1mm, maximum cross section of 45x45mm.
Maintaining the ring below the yielding.
The tests need to be always supervised.
Ring holder

## Slide 4: FEM simulation of jacket ring
20°
0.5mm
Cyclic symmetry with 20° sectors (only jacket)
Ring/holder contact
Imposed displacements to holder (0.1mm steps)
Yielding of 316LN @20°C is 220MPa (E=200GPa)*
*ITER Appendix A, Materials Design Limit Data (222RLN)
Holder
Jacket Ring
L=22.7mm
H=29.1mm
R=502.5mm

## Slide 5: FEM simulation of jacket ring
Maximum allowable radial displacement is 0.5mm
Estimated loading cycle: 0.1mm to 0.5mm.
 | Internal | External
σH max [MPa] | 193.98 | 192.97
σH min [MPa] | 191.10 | 172.79

## Slide 6: Fracture mechanics simulation
A code for the simulation of crack growth has been developed following the ITER codes*
The results are cross-checked (in terms of stress intensity factors) with FEA results
Input data:
  a=1mm; 2c=3mm
  t=3mm; w=29mm
  σt=200MPa
*ITER Magnet Structural Design Criteria Part 1: Main Structural Components and Welds (2FMHHS)

## Slide 7: Crack simulation on jacket ring: external crack
External crack
Initial defect area: 4mm2 (2mm2 with SFA=2*)
Semi-elliptical crack: a=1.3mm, 2c=3.91mm
σt=182.6MPa, σb=9.8MPa, R=0.2
Paris coefficient: C=3.86e-11, m=2.394 (316LN)
da = 0.001mm
KIC=86.7MPa (130MPa with SFK=1.5*)
*ITER Magnet Structural Design Criteria Part 1: Main Structural Components and Welds (2FMHHS)

## Slide 8: Crack simulation on jacket ring: external crack
Analytical vs Numerical:
Similar trend
Analytical model provides greater values (~8%) with respect to numerical ones
This discrepancy can be attributable to:
  the conservative nature of the analytical solution
  inherent structural stiffness of the ring geometry, which provides greater crack-closure constraint compared to the equivalent flat plate model

## Slide 9: Crack simulation on jacket ring: external crack
Number of cycles before failure: 26931 (with SFN=2)
Failure mode: a=0.8·t
K<<KIC
Minimum 25 days test
(20s/cycle - 6h/d - no issues)

## Slide 10: Crack simulation on jacket ring: Internal crack
External crack
Initial defect area: 4mm2 (2mm2 with SFA=2*)
Semi-elliptical crack: a=1.3mm, 2c=3.91mm
σt=191.6MPa, σb=0MPa, R=0.2
Paris coefficient: C=3.86e-11, m=2.394 (316LN)
da = 0.001mm
KIC=86.7MPa (130MPa with SFK=1.5*)
*ITER Magnet Structural Design Criteria Part 1: Main Structural Components and Welds (2FMHHS)

## Slide 11: Crack simulation on jacket ring: Internal crack
Analytical vs numerical:
Similar trend.
Analytical model provides greater values (~4%) with respect to numerical ones.
As for the external crack, the overestimation can be attributed to the conservative nature of the analytical solution and to the ring geometry.

## Slide 12: Crack simulation on jacket ring: Internal crack
Number of cycles before failure: 25119 (SFN=2)
Failure mode: a=0.8·t
K<<KIC
Minimum 23 days test
(20s/cycle - 6h/d - no issues)

## Slide 13: Summary
A FEM model of the jacket has been developed and used to evaluate the experimental loading cycle amplitude, considering all the boundary constraints.
A code the crack growth has been developed based on ITER design codes. The results, in terms of SIF along crack front, were successfully cross-checked against numerical simulations.
Once available, the experimental results will be compared to those obtained by the developed code.

## Slide 14: Advanced finite element modeling to support the experimental test MAG-S.01.04-T023-D001
F. Giorgetti
WPMAG 2025 Final Meeting, 15-Jan-2026

## Slide 15: Scope of the work
Compare the crack propagation, on metallic components, predicted by means ITER design codes and experimentally derivate.
A series of tests are planned to study the evolution of artificial defects on a steel jacket. Specifically, tests will be performed on a linear jacket (with a 4-point bending test) and on a ring jacket.
Ring jacket test will be performed at ENEA making use of the Testing Facility already developed to test the ITER pre compression rings.