## Slide 1: Purchase of a FIB-SEM equipped with EDS, EBSD and laser ablation chamber
Alessandro Rufoloni, Angelo Vannozzi
ENEA CR Frascati
Superconductivity Lab.
WPMAG Final Meeting – 16 Jan. 2026
MAG-T.01.01-T043-D001

## Slide 2: Aim and outline
In-situ preparation and characterisation of striated commercial REBCO tapes
FIB-SEM
REBCO tape sample description
Proposed tests
Results
Comparison
Conclusion

## Slide 3: FIB-SEM configuration
Possible solutions
FIB-SEM
Laser beam
+
FIB-SEM
Laser beam
FIB-SEM
Stand alone
Pre-chamber
Fully integrated

## Slide 4: REBCO coated conductor
4
Multilayer architecture deposited on Hastelloy metallic substrate
REBCO: REBa2Cu3O7-x (RE=Y or rare earth)
Non-symmetric sample, REBCO film on one side only
Region of interest: REBCO+buffer stack
Analysis of film integrity and interfaces
Copper (10 µm)
Silver (2 µm)
REBCO (2 µm)
Buffer stack (350 nm)
Hastelloy substrate (60 µm)
Typical tape width 4 mm (or 12 mm)
50–100 µm
Copper (10 µm)
Silver (1 µm)

## Slide 5: REBCO coated conductor filamentisation
5
Laser patterning to produce REBCO filaments
Typical strip width 0.4–1 mm
Useful strip length: > 4 mm
Hastelloy substrate (60 µm)
Copper (10 µm)
Silver (1 µm)
Cu
Ag
REBCO
Buffer
Typical tape width 4 mm (or 12 mm)
50–100 µm
Li et al. SuST 37 (2024)

## Slide 6: Analysis of striated REBCO coated conductor
Copper (10 µm)
Silver (1 µm)
6
Analysis of cross-sectional REBCO lamella extracted from patterned tape
Cu
Ag
REBCO
STEM for defect detection (cracks, voids, inhomogeneities) across filaments
EDS (linescan or map) to detect possible compositional inhomogeneity produced by laser
EBSD to detect possible phase modification or amorphisation produced by laser patterning
Hastelloy substrate (60 µm)
Buffer
Li et al. SuST 37 (2024)
Shanghai SST 2206 tape was selected for tests

## Slide 7: Experimental
Three main tests requested
Tape filamentisation by laser patterning to obtain several REBCO strips running parallel to the rolling direction
Extraction of cross-sectional lamella on patterned REBCO tape in correspondence of strip edge
Analysis of cross-sectional lamella with focus on possible damage caused by laser patterning, including STEM, EDS and EBSD

## Slide 8: 1) Stand-alone laser
REBCO tape striation
Size 3 channels 6.5 mm x 0.4 mm x 40 µm depth
Power 100 mW, Time 57’, f=20 kHz
Spot diam. 10 µm

## Slide 9: 1) Stand-alone laser
Cross section (filament edge)
Bulk cross section obtained with plasma-FIB
Ablated material redeposited on the sample
Cu
Pt
Ablated material
REBCO
Ag
Cu
Cu
Hastelloy

## Slide 10: 2) Integrated laser unit (pre-chamber)
REBCO tape striation
Size 20 mm x 35 µm x 23 µm (9 trenches)
Power 50 mW, pulse 350 fs, 10 passes, removal rate c. 2 µm/pass, strip width 350 µm
Time 2.5 h
inhomogeneous milling (bending effect)

## Slide 11: 2) Integrated laser unit (pre-chamber)
REBCO tape striation
Size 20 mm x 35 µm x 23 µm (9 trenches)
Power 50 mW, pulse 350 fs, 10 passes, removal rate c. 2 µm/pass
Time 2.5 h
Thick ablated material layer redeposited on the strips
inhomogeneous milling (bending effect)
Gap 35 µm
Strip 350 µm

## Slide 12: 2) Integrated laser unit (pre-chamber)
REBCO tape striation
Size 20 mm x 35 µm x 23 µm (9 trenches)
Power 50 mW, pulse 350 fs, 10 passes, removal rate c. 2 µm/pass
Time 2.5 h
Eventually measured I-V (whole sample) and VSM+OM (box) next presentations
inhomogeneous milling (bending effect)

## Slide 13: 2) Integrated laser unit (pre-chamber)
Cross section (filament edge)
STEM imaging on thin cross-sectional lamella (4 different signals)
Pt
Ablated material
REBCO
Ag
Cu
Hastelloy
REBCO
BL
2nd phases

## Slide 14: 3) Fully integrated laser beam
REBCO tape striation
Size 1 mm x 30 µm x 15 µm (6 trenches)
Power c. 1.2 W, pulse energy 15 µJ, 1-2 passes, removal rate c. 15 µm/pass
Overlap between meeting segments critical
After plasma FIB rocking polishing
W
REBCO
Ag
Cu
Hastelloy

## Slide 15: Additional analysis: EDS (producer 2)
EDS on extracted cross-sectional lamella (edge)
Contamination due to Ni ablated from the substrate
Slight thickness effect (lamella thickness inhomogeneity)
No obvious oxygen content gradient
Ba- L series
O- K series
Ag- L series
Eu- L series
Ni- K series
Ce-L series

## Slide 16: Additional analysis: EBSD (producer 2)
EBSD on extracted cross-sectional lamella (edge)
Majority of REBCO film c-axis oriented; a-axis regions possible, but further investigation required

## Slide 17: Comparison
Comparison of different configurations proposed
Producer | 1 | 2 | 3
FIB source | Xe | Ga | Xe
Laser | stand alone | integrated (pre-chamber) | integrated
Laser type | ps/fs | fs | fs
Single trench length (mm) | 45 | 40 | 1
Load lock | Y | Y | N

## Slide 18: Comparison
Comparison of different configurations proposed
Producer | 1 | 2 | 3
FIB source | Xe | Ga | Xe
Laser | stand alone | integrated (pre-chamber) | integrated
Laser type | ps/fs | fs | fs
Single trench length (mm) | 45 | 40 | 1
Load lock | Y | Y | N

## Slide 19: Conclusion
Preliminary test of REBCO tape filamentisation using laser beam integrated in FIB-SEM
Three FIB-SEM producers considered
Different FIB sources proposed: Xe (plasma FIB) or Ga (conventional source)
Three laser configurations proposed: stand alone, integrated (pre-chamber) and fully integrated
Single trench length varies considerably, from c. 1 mm to > 40 mm
After patterning a considerable amount of material ablated redeposited on the strips
EDS and EBSD on cross-sectional lamella feasible
Analysis of cross-sectional lamella to detect possible negative effects due to laser patterning revealed no apparent damage on REBCO film integrity

## Slide 20
Thank you
for your attention
angelo.vannozzi@enea.it