## Slide 1: Upgrade of Timing Counter ― Pixelated TC
PSI BVR MEG review meeting
Jan.14.2013, PSI
Yusuke Uchiyama

## Slide 2: Contents
Present performance
Concept of upgrade
R&D status
Schedule

## Slide 3: Present performance
Measurement of e+ time is limited by
  Uncertainty of propagation time from DC to TC: 75 ps
  TC single bar time resolution: 60 ps
    large TTS due to B-field
    position & incident-angle dependence
    long propagation to PMTs
  Time alignment among bars (calibration): 35 ps
  Electronics synchronization: 30 ps
Furthermore,
  PMTs are now working at the edge of proper operation point of hit rate (1 MHz)
  Not fit new DC in geometry (slanted PMTs, N2 bag)
76 ps

## Slide 4: Present performance
Measurement of e+ time is limited by
  Uncertainty of propagation time from DC to TC: 75 ps
  TC single bar time resolution: 60 ps
    large TTS due to B-field
    position & incident-angle dependence
    long propagation to PMTs
  Time alignment among bars (calibration): 35 ps
  Electronics synchronization: 30 ps
Furthermore,
  PMTs are now working at the edge of proper operation point of hit rate (1 MHz)
  Not fit new DC in geometry (slanted PMTs, N2 bag)
76 ps
Obviously, upgrade of TC is also necessary to overcome these.
Naturally, comes to the concept of…

## Slide 5: Concept of TC upgrade
Pixelated scintillator detector
  Small scintillator plates
  Several hundreds of counters
  Readout by SiPMs
Advantages
  Operational at high rate
    hit rate is not an issue
    less pileup, no double-hit by single e+
  Good single-counter resolution
    proved technology by PSI μSR group (smaller)
    less contributions from position and angle
    operational in B-field and in Helium gas
  Further improvement with multi-counter hits
    reduce electronics & calibration contributions as well

## Slide 6: Concept of TC upgrade
Pixelated scintillator detector
  Small scintillator plates
  Several hundreds of counters
  Readout by SiPMs
Advantages
  Operational at high rate
    Hit rate is no more an issue. Less pileup, no double-hit by a e+.
  Good single-counter resolution
    proved technology by PSI μSR group (but, for smaller counters)
    less contributions from position and angle
    operational in B-field and in Helium gas
  Flexible layout
  Further improvement with multi-counter hits
    reduce electronics & calibration contributions as well

## Slide 7: Single Counter
(k = area APD / area Scint. = 0.3)
Best result with PMT
→
(M.Moszynski, NIMA 337 (1993) 154.)
properties | BC418 | BC420 | BC422 | BC404
Light Output [% Anthracene] | 67 | 64 | 55 | 68
Rise Time [ns] | 0.5 | 0.5 | 0.35 | 0.7
Decay Time [ns] | 1.4 | 1.5 | 1.6 | 1.8
Wavelength [nm] | 391 | 391 | 370 | 408
Attenuation Length [cm] | 100 | 110 | 8 | 140
current TC bar
Ultra-fast scintillator + SiPMs
  High performance was already proved by PSI muSR group with smaller counter
    Better intrinsic resolution than PMTs
    Resolution scales as # of observed p.e.
    Small coverage can be recovered by using several SiPMs connected in parallel or series
Measured resolution for small counter
with SiPM readout
Established (basic) technology

## Slide 8: Single Counter
(k = area APD / area Scint. = 0.3)
Best result with PMT
→
(M.Moszynski, NIMA 337 (1993) 154.)
properties | BC418 | BC420 | BC422 | BC404
Light Output [% Anthracene] | 67 | 64 | 55 | 68
Rise Time [ns] | 0.5 | 0.5 | 0.35 | 0.7
Decay Time [ns] | 1.4 | 1.5 | 1.6 | 1.8
Wavelength [nm] | 391 | 391 | 370 | 408
Attenuation Length [cm] | 100 | 110 | 8 | 140
current TC bar
Ultra-fast scintillator + SiPMs
  High performance was already proved by PSI muSR group with smaller counter
    Better intrinsic resolution than PMTs
    Resolution scales as # of observed p.e.
    Small coverage can be recovered by using several SiPMs connected in parallel or series
Measured resolution for small counter
with SiPM readout
(≒BC422)
2 MPPCs
Series connection
A. Stoykov et al. NDIP 2011
MPPC: S10362-33-050C
(Hamamatsu, 3x3mm2, V~70V)
Established (basic) technology

## Slide 9: Multi-counter hit
MC simulation
5-counter hit

## Slide 10: Multi-counter hit
Multiplicity is studied with MC
Measured resolution is used
Overall resolution is significantly improved by combining multiple hits
〈Nhit〉=4.8
Overall time resolution vs Nhit
Weaken electronics (jitter, sync) and calibration contribution as well

## Slide 11: Expected performance
Single pixel counter
  Base design (30x90x5 mm3, BC422):50ps
Multiple hit
⇒ the average time resolution :
30–35 ps (63% ↓)
resolution
130 ps → 84 ps (35% ↓)
  Gamma side: 67 ps → 76 ps
  Track length: 75 ps → 11 ps
  Timing counter: 76 ps → 30–35 ps
104 ps → 35 ps

## Slide 12: Items for R&D
Prove the performance for our application
Larger single counter
Readout system (cabling, electronics…)
Multiple hit principle
Other possible issues
The number of electronics channels (~600 pixels×2 readout) and SiPMs (~600 pixels×6), and cost
Temperature coefficient of SiPM gain
Radiation hardness of SiPM

## Slide 13: R&D of single counter

## Slide 14: Prototype and Items
Basic study & optimal operation voltage
Scintillator type (fast rise time, longer attenuation,,,)
Dimension of a counter
# of SiPMs & manufacturer
Reflector

## Slide 15: Setup
3 or 4 MPPCs
Series connection
Test counter
  HAMAMATSU MPPC (S10362-33-050C)
    3x3 mm2, 50mm-3600 pixels
  glued with optical grease (OKEN6262A) to be able to detach
  Swap scintillators
90Sr b-source (<2.28 MeV)
Trigger with reference counter (RC)(5x5x5mm3, 1 MPPC)
  Measured time resolution : σRC = 28 ps (for Edep > 0.5 MeV)
Voltage preamp from PSI (Gain~20, 600 MHz bandwidth)
  Shaping with high-pass filter & pole-zero cancellation
Long cable before preamp
KEITHLEY Picoammeter for MPPC bias (HV)
DRS4 evaluation board (1 chip, 4 ch) @ 5 GSPS

## Slide 16: Analysis
Signal time is picked-off by Constant-Fraction method (~10%)
  very leading-edge is relevant to precise timing
e hit time is reconstructed by the average of times measured at the both ends
Resolution of test counter is evaluated from (t1 + t2)/2 – tref
trise~1.7 ns
baseline restoration by the
pole-zero cancellation @ preamp

## Slide 17: Bias dependence
Improvement in resolution as DV
then degraded because of higher dark noise
Following test done at DV=2.3V
BC422
6x3x0.5 cm3
3M mirror film
3 MPPCs
connected in series
Break-down voltage
operation

## Slide 18: Wrapping w/ reflector
We tested Aluminized Mylar (from Detector group), Teflon tape, 3M radiant mirror film (from Genoa group).
Teflon
3M
Al Mylar
  Setup
  BC422, 60x30x4.5 mm3
  3 MPPCs in series each side
  32 ns LEMO cable before preamp
  Preamp 1 kOhm
  Signal divided by two
  RC: 71.5 V, 2.23 uA, signal divided by 4
About all, position scanning was done.
non-adhesive side

## Slide 19: Results w/ wrapping
3M-film is the best for time resolution.
The non-uniformity may come from way of wrapping (?)
naked (no wrapping)
Al-Mylar
Teflon
3M mirror film
Amplitude (V)
Naked
Al-Mylar
3M film
Scan 15 points (1cm interval in X,Y)
Resolution maps

## Slide 20: Optimization
Tested different size
  The larger, the worse as expected from # p.e.
    longer → attenuated, wider → smaller coverage
    BC422 is always better than BC418
      fast rise-time has larger impact than long attenuation
    Degradation is not significant
      ~50 ps has been achieved for 9 cm long counter
  The larger, the higher efficiency and the higher multiplicity
    optimization should be done with respect to exp. sensitivity
  Not sensitive to thickness
    as far as using 3x3mm2 SiPMs # p.e is constant
    Thinner is better for multiplicity
4 MPPC connection improves
Optimization is still in progress
  ex) Manufacturer of SiPM (so far, only Hamamatsu)
BC422, 3 MPPCs
However

## Slide 21: Temperature dependence
Relatively large temperature coeff. of break-down voltage for MPPC (56mV/℃) might be an issue
  Temp. variation in COBRA is a few ℃
  Time shift due to this is expected to be ~15 ps
    doesn’t seem a big issue
  Possible solutions
    Improve detector temp. stabilization
    KETEK SiPM with lower temp. coeff. (<1%/℃)
~100ps/V
Measured time shift vs over voltage

## Slide 22: B-field test
Tested effect of B-field on SiPM
  1.25T applied to AdvanSiD SiPM
    both in orthogonal and parallel
  Measurements with Laser
    405nm, 100 ps width
  No appreciable effects observed in gain nor timing
Pavia NMR Lab
B-field up to 1.25 T
Put SiPMs
Amplitude
Over Voltage (V)
Time position (psec)
Time resolution (FWHM) (psec)
B-field effects (OFF, ON (orthogonal, parallel))

## Slide 23: Optimization study (MC)
Longer counter configurations show better overall resolution due to high multiplicity as well as higher efficiency
Final configuration is being designed with respect to exp. sensitivity
  MC study with the measured resolutions as inputs
    sensitivity by simplified (Punzi’s) figure of merit
  Trade off b/w performance vs cost
  Other factors for longer counter (hit rate, pileup) are under study
better
for fixed sensitivity
overall resolution
efficiency
for 3 cm width 3 MPPCs

## Slide 24: Progress post proposal submission
beam test at the end of RUN 2012 at pE5

## Slide 25: Introduction
We performed a beam test of the prototype of pixelated TC
  Started discussion after the collaboration meeting (30/Nov) in accordance with the proposal of beam test for Active TARget
  Decided on 12/Dec to carry out with (only) two counters and started preparation
  Successfully carried out tests with beam for several configurations on 20/Dec.
    parasitically to ATAR with less priority
  Ended with accelerator shutdown

## Slide 26: Motivations
Test with MIP
  Lab test has been done with 90Sr (up to 2 MeV)
    dE/dx is different from MIP
    Not penetrate two counters
Noise situation in the area
  even though not exactly same as final one in COBRA
High hit rate condition
See temperature variation and its impact
Find any possible issues towards the planned beam test in this spring

## Slide 27: Counters
2 BC422 counters
  C1: 6x3x0.5 cm3
  C2: 9x3x0.5 cm3
  3 MPPCs at each side connected in series
  Wrapped with 3M radiant mirror film
  Pt100 sensors are attached
  Light shielded by black tape
Tested in lab with 90Sr in advance
C1
C2
Expected resolution of two-counter measurement
37 ps (C1: 48, C2: 57 ps) for MIP
(no inter-counter jitter)

## Slide 28: Setup
DRS
pE5 area
Independent DAQ system
C1
C2
7.4 m

## Slide 30: Setup
z
y
C2
C1
BTS
ATAR
Collimator
m+ or e+ beam (28 MeV/c)
70 cm
air

## Slide 31: Configurations
Parallel | Cross | Staggered
C1+RC1 | C2+RC1 | C2+RC1+RC2
C1
C2
Trigger by 4 (or 3) coincidence

## Slide 32: Signal amplitude
Lower Edep for MIP than for 90Sr
Longer tail for Landau distribution
Energy deposit
― 90Sr (MC)
● 90Sr (data)
― 50MeV e+ (MC)

## Slide 33: Noise
Noise increased by ~40%
Picked up coherent noise on SiPM boards?
  RC with PCB (PSI) doesn’t show increase noise
Due to high hit rate?
Under investigation
Beam test
Lab
Correlation b/w counters

## Slide 34: Some analysis & results
Propagation time b/w two counters depends on particle’s direction
  Particles are not collimated.
    Many comes from off-target
    No tracker, no calorimeter
    Uncertainty of the counter’s alignment
  Corrected by reconstructed hit points
    only one direction
Position dependence is observed
  consistent with that in Lab data
  corrected with rec. position

## Slide 35: Parallel configuration
C1
5 mm
C2
w/o position cut, two-counter res = 40 ps
w/ position cut for 5mm, 36 ps
Deteriorated as noise
(T1 – T2) / 2 (sec)
Preliminary
Two-counter resolution is evaluated
Noise (C1-1) vs resolution

## Slide 36: Cross configuration
Measured two-counter resolution: 38 ps
Improved by selecting center (5x5 mm2):
  eliminate (reduce) uncertainty of path length
  34 ps
C1
ch4
ch3
ch1
ch2
View from top
Preliminary
Consistent with lab measurements
(T1 – T2) / 2 (sec)

## Slide 37: To do
Complete different SiPMs study (underway)
  Hamamatsu, KETEK, SensL, AdvanSiD
  Test new type sensors once we get samples
  Test temperature variation effect in thermostatic chamber
Prototype with several counters
  beam test in spring to prove multi-hits scheme
  make schedule
Electronics (WaveDREAM) test before beam test
  Can manage without preamp? ⇒ offline shaping
  Sampling speed (5G? 3G? 1.6G?) ⇔ trigger latency
Establish calibration methods
  Laser and/or Michel
Design full detector (support, cabling, etc.)

## Slide 38: Different SiPMs tests
We have been testing different SiPMs
  Breakdown voltage, its variation, stability
  Temp. sensitivity
  Gain, PDE
  Timing performance
Temp. coeff.
65 mV/℃
Temp. coeff.
40 mV/℃
AdvanSiD
BV~44 V
HAMAMATSU
BV~72 V

## Slide 39: New SiPMs
Recent developments of new types of SiPM
Large impact on time resolution is expected for our application, too
Hamamatsu MPPC

## Slide 40: Calibration
Important item for actual operation
Laser calibration
  Light pulse from a laser system
    Hamamatsu PLP10-040
      70 ps width, 405 nm wavelength
  Distribute via optical fibers
Michel (track-based)
  in-situ calibration using data itself
  Develop a technique similar to position alignment
    like Millipede or Linear-fit algorithms
    to calibrate all channels simultaneously
Finally, with RMD

## Slide 41: Further improvements, Backup solutions
Low-power high-bandwidth preamplifier
  For good S/N, even in case of large noise
  O(mW), 0.35–1 GHz
  Work in progress in Genoa
    achieved 0.9 ns rise-time with 8 mW
Study on organic scintillator
  higher light yield than plastic scint. (x3) with fast decay time (a few ns)
  ‘p-terphenyl’ (C18H14) is a good candidate
  Direct comparison with plastic scint. is under way with small counters.
90Sr
p-terphenyl crystal (19x11x3 mm3)
2 SiPMs

## Slide 42: Manpower & expertise
Italy (4.5) | Experience of present TC construction & operation. Scintillator. Calibration, analysis. Mechanics, electronics.
Japan (2) | SiPM. Analysis. Simulation, design. Students for intensive work & education.
USA (2) | Tracking (combined analysis with DC), calibration.
PSI (1) | Experience of muSR counters (scinti.+SiPM). Electronics, machining.
The new TC team

## Slide 43: Schedule
Milestones: two staged prototypes and beam tests
Several counters prototype
  Beam test in this spring
  @DAFNE Beam test facility (BTF) at Frascati
  ~30 counters
  Prove multi-hit scheme
~1/4 scale prototype
  Beam test in COBRA
  A few 100 counters
  Prove full performance in situ

## Slide 44: Conclusion
Upgraded timing counter
  pixelated TC with ultra-fast scint. & SiPMs
  Significant improvement in timing resolution with multi counters hit
R&D successfully in progress
  Systematic study on single counter with source in Lab
  already achieved good resolution
    43ps for 6cm, 50ps for 9cm long counter
  Carried out 1st beam test with two counters
    confirmed the results in lab
    obtained useful information for further study
      higher noise, Edep, incident position angle effects
Expected resolution of 30–35 ps significantly reduces BG together with the new drift chamber

## Slide 45: Back up

## Slide 46: Time line
 |  | Dec.12 | Dec.13 | Dec.14 | Dec.15 | Dec.16
 |  | Made decision |  | Counter test in lab | Counter test in lab | 
Dec.17 | Dec.18 | Dec.19 | Dec.20 | Dec.21 |  | 
Counter setup (Pre-amplifiers, thermometer, light shielding, building stands, etc.) | Setup in area DAQ setup Check with cosmic ray | Setup at beam line Trigger setup Beam test 18:40- Parallel conf | Beam test -4:00 15:50-16:15 Parallel conf 17:25-19:00 Cross conf | Beam test 2:50-3:40 w/ RC1 4:40-5:00 Staggered conf 5:10-6:30 (beam shutdown) w/ RC1+RC2 |  | 

## Slide 47: Cost estimation

## Slide 48: 3MTM radiant mirror film
Multi-layer polymeric film.

## Slide 49: Position scan
Position dependence of time measurement
25ps
40ps
BC422
EJ228
position scan w/ reference counter
1 cm
1 cm
There is a little position dependence.
However, we can correct it

## Slide 50: Position reconstruction
⇒ attenuation length
⇒ scintillation light speed
BC422
EJ228
Resolution
~8 mm
BC422
EJ228
↓
Since attenuation length is long, position reconstruction by charge does not work.
Resolution ~14.5 mm
current TC Bar
σ~15 mm (MC)
from time diff.
c.f.

## Slide 51: Preamplifier
Using preamplifier developed at PSI
Parameters were optimized for mSR counter
  signal is larger and faster than ours
Optimize for our application
  Gain
    ~25 (increased by a factor 2)
  Shaping
    High pass + pole-zero cancelation (CR-RC-CR)

## Slide 52
Adjust time constant by the potentiometer
  to flat baseline

## Slide 53: Pulse shape
200 Ω (attenuation 2)
0 Ω (attenuation 4)
400 Ω (attenuation 2)
600 Ω (attenuation 2)
800 Ω (attenuation 2)
1000 Ω (attenuation 2)
:
:
:
1680, 1700 Ω
Waveform did not change so much.

## Slide 54
Shaping of high-pass + pole-zero cancelation is definitively necessary
But performance doesn’t depend on detail value of parameter
  following tests done with 1 kW
Measurement at DV = 2 V

## Slide 55: MC simulation
GEANT4 (gem4)
  Full detector simulation
    optimization of configuration, geometry
    study efficiency, overall spectrometer performance,
  Simulation for R&D
    counter simulation
      to understand the device and test results
      optimize single counter design
    preceding study by PSI muSR group
      NIMA 696, 40–45 (2012)
MPPC sensitive area
Scint. photon tracking
(reduced by factor 100 for this display)
Status
Basic material & geometry
  Scintillator (BC418 or 422), rise time, decay time
Started scint. photon tracking
  attenuation, reflection (polishedteflonair) by GEANT4
MPPC detection
  wavelength-dependent PDE
Device simulation to understand diode response?

## Slide 56
Hits in analyzer