## Slide 1: NMOR R&D for TRIUMF neutron EDM experiment
J. Martin, W. Klassen, C. Cerasani,
R. Mammei, M. Lang, T. Andalib,
D. Ostapchuk, and C. Bidinosti
The University of Winnipeg

## Slide 2: Need for Precision Magnetometry
The EDM (d) term violates CP.
New sources of CP violation required in e.g. electroweak baryogenesis.
Precision goal δdstat = 10-25 e-cm / cycle.
Requires average B known to δB < 16 fT, operating field B = 1 μT.

## Slide 3
Major Subsystems: Magnetic Shielding
Active
Passive
  Mag shield Room
  precision shield inside
Magnetic Field Generation
B0, B1, shim coils
Magnetometers
Fluxgates
Giant Magnetoresistance sensors
Xe/Hg Co-magnetometer
SQUIDs
Optical (NMOR)
UCN
comag
Internal Passive Shield
Magnetically Shielded Room
B0,1
Magnetic overview: nEDM experiment at TRIUMF
B0 1 µT
magnetometers
Active shield coils

## Slide 4: Magnetometers on the Market
Hg-199, Xe-129 10 fT in 100 s
Fluxgates 10 pT/rtHz magnetic
GMI 10 pT/rtHz
SQUID 1 fT/rtHz cryogenic
Alkali Atom Magnetometers
Ls OPM (Mx) ~10 fT/rtHz RF
NMOR 2 fT/rtHz all-optical
SERF 160 aT/rtHz B = 0
We have been studying NMOR-based magnetometers at U. Winnipeg
(as well as Xe-129, Fluxgates, GMI, and SQUID’s but that’s not in this talk)

## Slide 5: Magneto-optical rotation (Faraday rotation)
Circular birefringence ( rotates plane of polarization
Medium
Circular components of linear polarization
Linear effect from Zeeman splitting
Doppler width ~ GHz
B (Gauss)
M = 0
M = +1
M = -1
F = 1
F’ = 0
= 40 mT
D. Jackson-Kimball

## Slide 6: Nonlinear Magneto-Optical Rotation (NMOR)
Nonlinear = light modifies the medium
Hole burning: Light depletes ground-state velocity distribution.
Coherent dark state: Linearly polarized light selection rule creates coherent dark state.
burned
M = 0
M = +1
M = -1
F = 1
F’ = 0

## Slide 7: NMOR
Medium becomes linearly dichroic
Axis of alignment rotates by Larmor precession in the magnetic field
Net result: again, circular birefringence.
GHz → Hz
400 G → 2 μG
40 mT → 200 pT
B (Gauss)
2 μG
= 200 pT

## Slide 8: Removing Limitation on the Dynamic Range: AM/FM NMOR
Axis of alignment in the vapour rotates at , γ = 4.7 kHz/μT (85Rb)
Stroboscopic effect: flash the laser light at
Displaces the resonance from zero field to B.
FM NMOR
@ 500 Hz
Rb-87
D.F. Jackson-Kimball, et al.,
J. Appl. Phys. 106, 063113 (2009)

## Slide 9: Previous Work of Others
D. Budker, et al.,
PRL 81, 5788 (1998)
2.8 μG=
280 pT
D.F. Jackson-Kimball, et al.,
J. Appl. Phys. 106, 063113 (2009)
15 μG=
1.5 nT
=2 fT/rtHz
(T=35oC)
Zero field
Field ~ 500 uG
No modulation
FM NMOR

## Slide 10: Our Work - Apparatus

## Slide 11: Results from 2013-14 school year
Heat gunning campaign
New magnetic shield
C. Cerasani,
Hons. Thesis UW (2014)
NMOR results
@ zero field
Direct Measurement of Relaxation of Axis of Alignment
Clear indicator of issues with cell/coating quality (paraffin)
NMOR
Width
~130 uG
Graf, et al., Phys. Rev. A 72, 023401 (2005)
Us: τ = 2.1 ms = 1/γrel
Graf: τ = 0.4 s = 1/γrel
Cell from Precision Glassblowing, not designed with narrow capillary.

## Slide 12: Spring 2014, new cell (borrowed from D. Budker)
“Dear Dima and Brian,
Your cell works great!”
NMOR width ~ 1.5 μG = 150 pT
γrel ~ 1/0.36 s = 2.8 Hz
Consistent with Budker’s best results!
New cells from Balabas expected.
Possibly Precision Glassblowing
5 cm
Relaxation time
measurement

## Slide 13: Measuring small fields
± 46 fT steps in internal magnetic field supplied by calibrated coil.
Probe power 13 uW.

## Slide 14: Application: Large axial shielding factors measured with small applied fields
Wolfgang installing external coils.
Shielding factor = 1.3 x 107
Sensible variation with external coil spacing.
Paper in prep.
±1.45 uT applied externally
±475 pT applied internally
(calibration)
Optical rotation sees shielded field and calibration field

## Slide 15: Bandwidth, Noise, Linearity of Zero-Field Magnetometer
-3dB @ 2.5 Hz
30 fT/rtHz @ 1 Hz
  Room temp. cell and Rb
Linear
Also looked at Allan deviation
Sine wave internal field
Noise measurement:
All coils disconnected
Coils at various stable
currents

## Slide 16: AM NMOR
Width ~ 1 nT Consistent with Jackson-Kimball.
Larmor frequency consistent with fluxgate calibration 0.21 μT
Noise ~ pT
Still working to make it better.
Room temperature.
Quantifying existing magnetic noise, field quality.
ωmod = 2 kHz = 2 ωL

## Slide 17: Conclusions
System developed at U. Winnipeg approaches goal of fT-level magnetometry:
    NMOR widths consistent with best published values.
    ~30 fT/rtHz noise @ 1 Hz for zero field
    ~pT/rtHz noise @ 1 Hz for 0.22 uT applied field – we are working hard to understand and make it better.
Method applied to measurements of large axial magnetic shielding factors with small applied fields.

## Slide 18: Future
Make the best magnetometer we can based on this system and use it to study magnetic shielding, generation of magnetic fields.
Using this system will make us experts capable of employing a Cs-based fiberized system for the eventual nEDM experiment.
New: B. Patton et al “All-Optical Vector Atomic Magnetometer” PRL 113, 013001 (2014).
Thank you to D. Budker, B. Patton, and M. Solarz for the use of the cell.
Thanks NSERC, CRC, CFI, MRIF.