## Slide 1: The Single-Shot X-ray Spectrometer at SwissFel (PSSS)
Yunieski Arbelo, Benedikt Rösner, Christian David, Jens Rehanek, Christopher Arrell, Rafael Abela, Luc Patthey and Pavle Juranić :: Paul Scherrer Institut

## Slide 2: Shot-to-shot SASE spectra fluctuations
non-destructive shot-to-shot monitoring required
variations on intensity and photon energy!
B.D. Patterson et al. Physical Chemistry Chemical Physics, 12 21(2010)
“SwissFEL spectra measured with the PSSS!”

## Slide 3: The SwissFEL Single-Shot Spectrometer
“typical spectrum”
Rehanek, J., et al. Journal of Instrumentation 12.05 (2017): P05024.
M. Makita, et al. Optica 912-916 (2015)
Integrated | Aramis beamline (front-end)
Energy range | 4 to 13 keV
Resolution (∆E/E) | 2-5x10-5
Bandwidth | 0.5 %
Operation | Non-destructive (trans. >90%)

## Slide 4: Establishing parameters during commissioning
Calibrated shot-to-shot spectra retrieved!
“enabled energy ranges (crystals)”
“Expert Panel Set Up”
A parameters table implemented!
“enabled energy ranges (gratings)”
“4 to 12 keV”
“4 to 12 keV”

## Slide 5: Energy calibration of the PSSS
Dev. from Ref. XAS <1%
non-calibrated!
XAFS Spectra Library (http://cars.uchicago.edu)
http://exafsmaterials.com

## Slide 6: PSSS experimental resolution
Theor. Res. ∆E/E=2-5x10-5
Rehanek, J., et al. Journal of Instrumentation 12.05 (2017): P05024.
Exp. Res
FWHM≈0.48 eV

## Slide 7: Profile dependence on spectral intensity
“measured pulse profile”
“max. dev 3%”

## Slide 8: Limitations
Super important that the position and direction of the beam is preserved. Every 10 mgrad is about 5 eV offset. Bragg angles are very sensitive!
It cannot give large bandwidth. The crystals cannot bend that far.
Very sensitive to beam position. A 50 micron difference in vertical beam position is equivalent to about 0.3 degree offset in Bragg for the 200 radius mm crystal. That would be about 60 eV. Alignment is key!
Current radii between 75 and 200 mm give 0.5% FOV. Need 4 times smaller to get 2% FOV. Not possible.

## Slide 9: Beam position is important
The spectral intensity and clarity heavily depend on the beam position as well, as the graphs above show. Preserving the beam position is key to making sure both the spectral intensity and spectral energy are delivered well.

## Slide 10: Limitations
That means, when I get a call telling me that the PSSS got misaligned, and I check the slow position measurements on the PBPG and I see this:
The PSSS is not misaligned, and the beam moved. Which means the calibrations we did to make the simple user interface work no longer work.

## Slide 11: Conclusions
Shot-to-shot spectra were successfully acquired with the PSSS (non-destructive mode!).
XAS measurements enabled an accurate energy calibration of the spectrometer.
A resolution of ∆E/E ≈ 5x10−5 was experimentally determined for the PSSS.
User-friendly operation mostly implemented.
(Future: implementation as diagnostic tool for machine optimization)

## Slide 12: Performance optimization
factor x7
reduction
Kapton shielding
dKapton= 14 µm
TAir=<1 % at 4 keV
THe≈70 % at 4 keV
L=800 mm
dKapton= 50 µm

## Slide 13: Final work: GUI and processing
Creating a simplified user interface for ‘casual’ users.
Post-processing of data to get rid of noise, give values like the bandwidth and center of the spectrum.
Training—Nicole mentioned wanting to have a course for the operators once it is all polished.

## Slide 14: Wir schaffen Wissen – heute für morgen
Thank you very much for your attention!